Supplementary Table-I. LA-SF-ICP-MS U-Th-Pb dating methodology CAF, Stellenbosch   University

| Laboratory & Sample Preparation |  |
| --- | --- |
| Laboratory name | Central Analytical Facility, Stellenbosch University |
| Sample type / mineral | Igneous zircons |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, 1 µm polish to finish |
| Imaging | CL, LEO 1430 VP, 10 nA, 15 mm working distance |
| Laser ablation system |  |
| Make, Model & type | ESI/New Wave Research, UP213, Nd:YAG |
| Ablation cell & volume | Custom build low volume cell, volume ca.3 cm3 |
| Laser wavelength | 213 nm |
| Pulse width | 3 ns |
| Fluence | 2.5 J/cm-2 |
| Repetition rate | 10 Hz |
| Spot size | 30 µm |
| Sampling mode / pattern | 30 µm single spot analyses |
| Carrier gas | 100% He, Ar make-up gas combined using a T-connector close to sample cell |
| Pre-ablation laser warm-up (background collection) | 40 seconds |
| Ablation duration | 20 seconds |
| Wash-out delay | 30 seconds |
| Cell carrier gas flow | 0.3 l/min He |
| ICP-MS Instrument |  |
| Make, Model & type | Thermo Finnigan Element2 single collector HR-SF-ICP-MS |
| Sample introduction | Via conventional tubing |
| RF power | 1100 W |
| Make-up gas flow | 1.0 l/min Ar |
| Detection system | Single collector secondary electron multiplier |
| Masses measured | 202, 204, 206, 207, 208, 232, 233, 235, 238 |
| Integration time per peak | 4 ms |
| Total integration time per reading | 1 sec (represents the time resolution of the data) |
| Sensitvity | 20000 cps/ppm Pb |
| Dead time | 6 ns |
| Data Processing |  |
| Gas blank | 40 second on-peak |
| Calibration strategy | GJ-1 used as primary reference material, Plešovice & M127 used as secondary reference material (Quality Control) |
| Reference Material info | M127 (Nasdala et al. 2008); Plešovice (Slama et al. 2008) GJ-1 (Jackson et al. 2004) |
| Data processing package used / Correction for LIEF | In-house spreadsheet data processing using intercept method for LIEF correction |
| Mass discrimination | Standard-sample bracketing with 207Pb/206Pb and 206Pb/238U normalized to reference material GJ-1 |
| Common-Pb correction, composition and uncertainty | 204-method, Stacey & Kramers (1975) composition at the projected age of the mineral, 5% uncertainty assigned |
| Uncertainty level & propagation | Ages are quoted at 2 sigma absolute, propagation is by quadratic addition. Reproducibility and age uncertainty of reference material and common-Pb composition uncertainty are propagated. |
| Quality control / Validation | Plešovice: Wtd ave 206Pb/238U age = 338±4 (2SD, MSWD = 0.1) M127: Wtd ave 206Pb/238U age = 529±5 (2SD, MSWD = 0.8) |
| Other information | For detailed method description see Frei & Gerdes (2009) |