## Table S1 Detector Drift
Analyses of T1-G and K5 over extended time periods.
Elemental abundances (wt%)
| Sample | Analysis # | t after start (h) | V | Cr | Co | Ni | Cu | Zn | Ga | Rb | Sr | Y | Zr | Nb | Ba | La | Ce | Ta | Th | Na2O | MgO | Al2O3 | SiO2 | P2O5 | K2O | CaO | TiO2 | Fe2O3 | MnO
| T1-G | 1 | 3 | 0.010308 | 0.0059879 | 0.033014 | 0.00081338 | 0.0029391 | 0.010897 | 0.0047193 | 0.0071297 | 0.029638 | 0.0029753 | 0.029658 | 0.0055626 | 0.023842 | 0 | 0.01462 | 0.0099535 | 0.013305 | 2.3807 | 3.566 | 17.8971 | 57.1041 | 0.2761 | 2.016 | 8.2426 | 0.7925 | 7.3725 | 0.14693
| T1-G | 2 | 3.5 | 0.015119 | 0.0046888 | 0.033409 | 0.00038308 | 0.0026639 | 0.011166 | 0.0044401 | 0.0078223 | 0.029473 | 0.0023812 | 0.029579 | 0.0055996 | 0.025678 | 0 | 0.022957 | 0.010661 | 0.015057 | 1.9298 | 3.5479 | 17.8883 | 57.3943 | 0.32448 | 2.0083 | 8.3037 | 0.7966 | 7.4381 | 0.14736
| T1-G | 3 | 4 | 0.012925 | 0.0050204 | 0.032231 | 0.00080138 | 0.002679 | 0.011559 | 0.0039858 | 0.0078397 | 0.029442 | 0.0029671 | 0.029206 | 0.0055531 | 0.033551 | 0 | 0.017146 | 0.011361 | 0.014081 | 2.4271 | 3.5764 | 17.8058 | 57.1614 | 0.27388 | 2.0019 | 8.2415 | 0.79109 | 7.3557 | 0.14489
| T1-G | 4 | 4.5 | 0.011559 | 0.0050472 | 0.03346 | 0.00049368 | 0.0024685 | 0.011391 | 0.0048115 | 0.0066115 | 0.029761 | 0.0027464 | 0.0295 | 0.0047956 | 0.019352 | 0 | 0.024305 | 0.011053 | 0.013823 | 2.6073 | 3.4548 | 17.8606 | 57.11 | 0.20923 | 2.0093 | 8.2363 | 0.79609 | 7.3589 | 0.14622
| T1-G | 5 | 5 | 0.012852 | 0.0051453 | 0.032864 | 0 | 0.0020106 | 0.010898 | 0.0036238 | 0.007234 | 0.029226 | 0.0022453 | 0.029975 | 0.0055379 | 0.022069 | 0 | 0.020633 | 0.0099753 | 0.014499 | 2.0364 | 3.5295 | 17.8731 | 57.3568 | 0.30735 | 2.0268 | 8.2914 | 0.80289 | 7.4211 | 0.14585
| T1-G | 6 | 5.5 | 0.011477 | 0.0055741 | 0.032435 | 0.0005818 | 0.0025798 | 0.010728 | 0.0046105 | 0.0079603 | 0.029788 | 0.0019002 | 0.029619 | 0.0050729 | 0.020854 | 0 | 0.018582 | 0.010032 | 0.013706 | 2.3554 | 3.5883 | 17.8258 | 57.1948 | 0.23956 | 2.0097 | 8.2562 | 0.79468 | 7.3847 | 0.14542
| T1-G | 7 | 6 | 0.010684 | 0.0056161 | 0.032315 | 0.00044694 | 0.0032266 | 0.010891 | 0.0049439 | 0.0074219 | 0.029161 | 0.0022594 | 0.031384 | 0.0055431 | 0.027778 | 0 | 0.013836 | 0.011115 | 0.014407 | 2.3935 | 3.4524 | 17.8348 | 57.1981 | 0.31263 | 1.9956 | 8.273 | 0.78841 | 7.395 | 0.14539
| T1-G | 8 | 6.5 | 0.012302 | 0.0041243 | 0.033268 | 0.0002535 | 0.0021368 | 0.010667 | 0.0050034 | 0.0075881 | 0.029217 | 0.0018838 | 0.030093 | 0.0059303 | 0.021933 | 0 | 0.018755 | 0.010368 | 0.015032 | 2.1926 | 3.4472 | 17.854 | 57.3501 | 0.27659 | 2.018 | 8.2839 | 0.79207 | 7.4319 | 0.14517
| T1-G | 9 | 7 | 0.012848 | 0.0040044 | 0.033003 | 0.00061427 | 0.0023637 | 0.011415 | 0.0046834 | 0.0080096 | 0.030241 | 0.0021422 | 0.03051 | 0.0057929 | 0.023146 | 0 | 0.02184 | 0.011836 | 0.014473 | 2.3562 | 3.5184 | 17.8775 | 57.1588 | 0.25896 | 2.0106 | 8.2647 | 0.79307 | 7.3986 | 0.14612
| T1-G | 10 | 7.5 | 0.016001 | 0.0043273 | 0.032368 | 0.000683 | 0.002438 | 0.011231 | 0.0044002 | 0.0073986 | 0.030503 | 0.0030613 | 0.030606 | 0.0053582 | 0.029593 | 0 | 0.021328 | 0.011509 | 0.015853 | 2.2879 | 3.525 | 17.8404 | 57.2118 | 0.26253 | 2.0062 | 8.2657 | 0.79774 | 7.4285 | 0.14753
| T1-G | 11 | 8 | 0.012016 | 0.0044383 | 0.032822 | 0.0007381 | 0.002287 | 0.011432 | 0.0049068 | 0.0081496 | 0.030293 | 0.0025046 | 0.030735 | 0.0056998 | 0.027531 | 0 | 0.017928 | 0.011099 | 0.015599 | 2.0932 | 3.4054 | 17.9363 | 57.3302 | 0.30039 | 2.0136 | 8.2951 | 0.80277 | 7.4569 | 0.14805
| T1-G | 12 | 8.5 | 0.011976 | 0.0050409 | 0.033189 | 0.00061019 | 0.0029962 | 0.011535 | 0.0045826 | 0.0082401 | 0.030486 | 0.0023153 | 0.029558 | 0.0065915 | 0.025488 | 0 | 0.022596 | 0.011893 | 0.016587 | 2.4381 | 3.68 | 17.8487 | 56.997 | 0.2335 | 1.9987 | 8.2419 | 0.80168 | 7.3905 | 0.14625
| T1-G | 13 | 9 | 0.014661 | 0.0049531 | 0.033478 | 0.000323 | 0.0027227 | 0.011065 | 0.0039642 | 0.0074119 | 0.030648 | 0.0018361 | 0.030519 | 0.0056558 | 0.024218 | 0 | 0.023933 | 0.012628 | 0.014349 | 2.4223 | 3.5465 | 17.7899 | 57.0976 | 0.26879 | 2.0118 | 8.2834 | 0.79302 | 7.4169 | 0.14734
| T1-G | 14 | 9.5 | 0.018 | 0.0052606 | 0.032372 | 0 | 0.0023961 | 0.010851 | 0.0044396 | 0.0082661 | 0.030439 | 0.0023845 | 0.029652 | 0.0052234 | 0.02707 | 0 | 0.025554 | 0.012333 | 0.015008 | 2.7382 | 3.5028 | 17.7797 | 56.9289 | 0.27692 | 1.999 | 8.2288 | 0.79739 | 7.372 | 0.14721
| T1-G | 15 | 10 | 0.014559 | 0.0051758 | 0.033217 | 0.00023016 | 0.0027823 | 0.011009 | 0.0039765 | 0.0073907 | 0.030217 | 0.0026831 | 0.030615 | 0.0057505 | 0.026398 | 0 | 0.022602 | 0.010097 | 0.016116 | 2.4107 | 3.5064 | 17.8502 | 57.1648 | 0.25921 | 1.9936 | 8.254 | 0.79025 | 7.4007 | 0.14731
| T1-G | 16 | 10.5 | 0.015543 | 0.0053828 | 0.032776 | 0.00018792 | 0.0025079 | 0.011012 | 0.0040571 | 0.0079978 | 0.030235 | 0.0024228 | 0.030769 | 0.0068138 | 0.021505 | 0 | 0.025246 | 0.011581 | 0.014375 | 2.3308 | 3.48 | 17.8675 | 57.2005 | 0.25466 | 2.0114 | 8.2685 | 0.80148 | 7.4167 | 0.146
| T1-G | 17 | 11 | 0.012262 | 0.004888 | 0.033204 | 0.00062879 | 0.0025137 | 0.011387 | 0.0043616 | 0.006815 | 0.029871 | 0.0027882 | 0.03042 | 0.005771 | 0.01822 | 0 | 0.020376 | 0.012094 | 0.014739 | 2.3654 | 3.5743 | 17.6582 | 57.2254 | 0.26177 | 2.0173 | 8.2973 | 0.80784 | 7.4344 | 0.14776
| T1-G | 18 | 11.5 | 0.014126 | 0.006202 | 0.032279 | 9.9942e-05 | 0.0024395 | 0.011412 | 0.0047323 | 0.007522 | 0.030087 | 0.002253 | 0.030465 | 0.0059132 | 0.021073 | 0 | 0.025756 | 0.01029 | 0.01427 | 2.0155 | 3.489 | 17.834 | 57.3578 | 0.30085 | 2.0174 | 8.3356 | 0.80815 | 7.4767 | 0.14611
| T1-G | 19 | 12 | 0.016478 | 0.0055096 | 0.033183 | 8.7712e-05 | 0.0026267 | 0.011359 | 0.0042601 | 0.0081106 | 0.030571 | 0.0021454 | 0.030532 | 0.0062715 | 0.029103 | 0 | 0.023654 | 0.011787 | 0.015865 | 2.3882 | 3.4751 | 17.8821 | 57.116 | 0.26085 | 2.0078 | 8.282 | 0.79249 | 7.416 | 0.14786
| T1-G | 20 | 12.5 | 0.013166 | 0.0061096 | 0.033285 | 0.00011084 | 0.003369 | 0.011081 | 0.004172 | 0.0083077 | 0.029421 | 0.002261 | 0.031331 | 0.0060458 | 0.016707 | 0 | 0.024761 | 0.01028 | 0.014472 | 2.0517 | 3.5615 | 17.849 | 57.2691 | 0.28133 | 2.0156 | 8.3163 | 0.81143 | 7.4815 | 0.14765
| T1-G | 21 | 13 | 0.012483 | 0.0048527 | 0.032854 | 0.00013678 | 0.002266 | 0.010875 | 0.0044471 | 0.0073232 | 0.029836 | 0.0020722 | 0.030085 | 0.0060857 | 0.022531 | 0 | 0.020134 | 0.011231 | 0.014194 | 2.1438 | 3.5323 | 17.9774 | 57.1907 | 0.28095 | 2.008 | 8.2739 | 0.80058 | 7.4337 | 0.14725
| T1-G | 22 | 13.5 | 0.015669 | 0.0053036 | 0.032481 | 5.3951e-05 | 0.0027279 | 0.011361 | 0.0045836 | 0.007315 | 0.029851 | 0.0024607 | 0.031157 | 0.0057099 | 0.019789 | 0 | 0.025104 | 0.012039 | 0.015512 | 2.5112 | 3.5685 | 17.8035 | 57.005 | 0.2478 | 2.0072 | 8.2674 | 0.80127 | 7.4201 | 0.14679
Quantitative µ-XRF spot analyses for T1-G (left axis, white diamonds, measurements started 2013-04-20) and K5 (right axis, black circles, measurements started 2013-05-13) over 2.5 days. Black lines are lines of best fit for the T1-G data. T1-G analyses were carried out every 30 minutes after a single spectrometer calibration at the beginning of the experiment and some elements show small, systematic variations over time, indicating spectrometer drift. K5 analyses were carried out in batches with spectrometer calibration carried out at the beginning of each batch and show that regular recalibration reduces drift in the data. This figure shows how the data vary over time for four elements (Co, Sr, Fe2O3 and SiO2) which are representative of the behaviour of all analysed elements. The T1-G data shows that, even without regular calibration, most elements were stable or showed barely detectable changes over this time period (see SiO2). Four elements (Co, Ba, Ce and Ti – see Co on Fig.) showed a significant linear change over time (i.e. the drift exceeded the scatter in the data). The characteristic x-rays from these elements overlap with other elemental peaks and the observed drift is likely due to increasingly inaccurate peak deconvolution with time since energy calibration. A further four elements (V, Cr, Sr and Zr – see Sr on Fig.) showed an obvious but insignificant linear change over time of ~1.5% (i.e. the scatter exceeds the drift). Three elements (CaO, Fe2O3, MnO – see Fe2O3 on Fig.) show periodical drifts over time. These observed changes are small and only amount to a few tens of ppm for trace elements and hundreds of ppm for major elements. The drift becomes particularly noticeable ~15 hours after the spectrometer calibration. The K5 data shows that regular recalibration of the spectrometer using the pure Cu standard reduces this drift and increases the reproducibility of the results. | T1-G | 23 | 14 | 0.013268 | 0.0051147 | 0.033447 | 0.00041402 | 0.0031265 | 0.011199 | 0.0047941 | 0.0085907 | 0.030515 | 0.0023808 | 0.029975 | 0.006006 | 0.021731 | 0 | 0.022487 | 0.011522 | 0.014739 | 2.2868 | 3.5551 | 17.8757 | 57.095 | 0.26843 | 2.0192 | 8.2816 | 0.80107 | 7.4505 | 0.14726
| T1-G | 24 | 14.5 | 0.013051 | 0.0047497 | 0.033356 | 0.00015081 | 0.002743 | 0.011346 | 0.0049299 | 0.0078367 | 0.030451 | 0.0022363 | 0.030779 | 0.0054121 | 0.022796 | 0 | 0.022238 | 0.011529 | 0.013618 | 2.3105 | 3.5003 | 17.7353 | 57.1722 | 0.31078 | 2.0103 | 8.3338 | 0.79364 | 7.4681 | 0.14779
| T1-G | 25 | 15 | 0.016849 | 0.0057217 | 0.031679 | 0.0010141 | 0.002443 | 0.010765 | 0.0045396 | 0.0080526 | 0.030585 | 0.0028406 | 0.030331 | 0.0056853 | 0.013896 | 0 | 0.02788 | 0.010986 | 0.01392 | 2.4897 | 3.6144 | 17.8633 | 56.952 | 0.28305 | 2.0032 | 8.2348 | 0.79694 | 7.3989 | 0.14639
| T1-G | 26 | 15.5 | 0.01571 | 0.0058587 | 0.03209 | 0.00060578 | 0.0023885 | 0.011509 | 0.0043343 | 0.0080415 | 0.03104 | 0.0029008 | 0.03086 | 0.0052291 | 0.020808 | 0 | 0.026759 | 0.010851 | 0.014029 | 2.2189 | 3.424 | 17.8963 | 57.2139 | 0.25669 | 2.015 | 8.3151 | 0.80551 | 7.4823 | 0.14912
| T1-G | 27 | 16 | 0.012408 | 0.0060988 | 0.032153 | 2.987e-05 | 0.002394 | 0.011293 | 0.0046276 | 0.007737 | 0.030359 | 0.0024525 | 0.030407 | 0.0056155 | 0.016725 | 0 | 0.027212 | 0.010318 | 0.014911 | 1.9444 | 3.4773 | 17.9069 | 57.4115 | 0.2722 | 2.0154 | 8.3204 | 0.80478 | 7.4854 | 0.14697
| T1-G | 28 | 16.5 | 0.013345 | 0.0055226 | 0.031235 | 0.00038715 | 0.0023001 | 0.011078 | 0.0046267 | 0.0074858 | 0.030237 | 0.0024452 | 0.029685 | 0.0057604 | 0.017066 | 0 | 0.026022 | 0.010597 | 0.015941 | 2.8389 | 3.5854 | 17.697 | 56.9171 | 0.25175 | 2.005 | 8.2161 | 0.79302 | 7.3373 | 0.1446
| T1-G | 29 | 17 | 0.014727 | 0.0055566 | 0.032075 | 4.2855e-05 | 0.0023735 | 0.010804 | 0.0048479 | 0.007515 | 0.030776 | 0.0027127 | 0.030691 | 0.0054391 | 0.025301 | 0 | 0.025562 | 0.0093488 | 0.014935 | 2.3443 | 3.5573 | 17.7151 | 57.1193 | 0.28197 | 2.0314 | 8.3136 | 0.7978 | 7.4681 | 0.14845
| T1-G | 30 | 17.5 | 0.014 | 0.0050856 | 0.032261 | 0.00054283 | 0.0027223 | 0.011335 | 0.0046072 | 0.0073062 | 0.029571 | 0.003061 | 0.02974 | 0.0055581 | 0.021408 | 0 | 0.025258 | 0.01116 | 0.014827 | 2.7348 | 3.4594 | 17.722 | 56.9892 | 0.30944 | 1.9991 | 8.2539 | 0.79289 | 7.3743 | 0.14655
| T1-G | 31 | 18 | 0.016234 | 0.0054019 | 0.031464 | 0.00012786 | 0.0030482 | 0.010983 | 0.0046584 | 0.0077818 | 0.029516 | 0.0019069 | 0.029052 | 0.0055992 | 0.017335 | 0 | 0.029467 | 0.009842 | 0.01401 | 2.0575 | 3.6664 | 17.8487 | 57.2446 | 0.26741 | 2.0152 | 8.3039 | 0.80454 | 7.4283 | 0.14699
| T1-G | 32 | 18.5 | 0.017733 | 0.005094 | 0.032213 | 0.00061969 | 0.0028918 | 0.011182 | 0.0044441 | 0.006889 | 0.02941 | 0.0023007 | 0.031016 | 0.0050692 | 0.024411 | 0 | 0.027785 | 0.011168 | 0.013844 | 2.1871 | 3.5652 | 17.8884 | 57.1655 | 0.23787 | 2.0237 | 8.2921 | 0.80463 | 7.4607 | 0.14876
| T1-G | 33 | 19 | 0.015712 | 0.0064415 | 0.032902 | 0.00029206 | 0.0031514 | 0.011728 | 0.0047713 | 0.0080744 | 0.03038 | 0.0027982 | 0.030478 | 0.0060308 | 0.019857 | 0 | 0.029406 | 0.010411 | 0.014266 | 2.0674 | 3.4854 | 17.8809 | 57.3172 | 0.26705 | 2.0122 | 8.3136 | 0.80574 | 7.4771 | 0.14668
| T1-G | 34 | 19.5 | 0.015993 | 0.0049598 | 0.031524 | 0.00037347 | 0.002349 | 0.010388 | 0.0039816 | 0.0082165 | 0.030832 | 0.0027887 | 0.031198 | 0.0054091 | 0.02106 | 0 | 0.028071 | 0.012175 | 0.015107 | 2.301 | 3.4445 | 17.8348 | 57.1885 | 0.26487 | 2.0136 | 8.3016 | 0.80877 | 7.4701 | 0.14781
| T1-G | 35 | 20 | 0.012151 | 0.0063824 | 0.03242 | 0.00016942 | 0.0027729 | 0.011103 | 0.0040942 | 0.0072648 | 0.030181 | 0.0024412 | 0.029898 | 0.0056879 | 0.014676 | 0 | 0.026279 | 0.011443 | 0.012765 | 2.3662 | 3.5248 | 17.8497 | 57.0898 | 0.24981 | 2.0161 | 8.2957 | 0.80458 | 7.4454 | 0.14803
| T1-G | 36 | 20.5 | 0.014399 | 0.0053942 | 0.032171 | 0.00017862 | 0.0021352 | 0.011305 | 0.0042998 | 0.0074086 | 0.030076 | 0.0018508 | 0.029388 | 0.0054361 | 0.016692 | 0 | 0.024857 | 0.011914 | 0.01317 | 2.5985 | 3.567 | 17.7621 | 57.0107 | 0.26233 | 2.0035 | 8.2496 | 0.80041 | 7.389 | 0.14622
| T1-G | 37 | 21 | 0.017532 | 0.0051796 | 0.031636 | 0.00030859 | 0.0034477 | 0.011309 | 0.0040708 | 0.009027 | 0.030313 | 0.002458 | 0.03115 | 0.0062598 | 0.031114 | 0 | 0.025833 | 0.011682 | 0.014309 | 1.9661 | 3.4316 | 17.8878 | 57.3806 | 0.26161 | 2.0156 | 8.3487 | 0.80415 | 7.5181 | 0.15002
| T1-G | 38 | 21.5 | 0.019137 | 0.0043124 | 0.031934 | 0.00040083 | 0.0028482 | 0.011087 | 0.004354 | 0.0075225 | 0.030943 | 0.0030368 | 0.031526 | 0.0059476 | 0.028465 | 0 | 0.025785 | 0.012069 | 0.015675 | 2.3147 | 3.502 | 17.7372 | 57.2072 | 0.24222 | 2.0117 | 8.3172 | 0.80114 | 7.482 | 0.14937
| T1-G | 39 | 22 | 0.016584 | 0.0055832 | 0.03265 | 0.00080021 | 0.0018665 | 0.011062 | 0.0039003 | 0.0073361 | 0.029935 | 0.0025471 | 0.02998 | 0.0054793 | 0.028025 | 0 | 0.020086 | 0.013119 | 0.013878 | 2.3969 | 3.4562 | 17.8619 | 57.0605 | 0.29293 | 2.0295 | 8.2807 | 0.79974 | 7.4525 | 0.14641
| T1-G | 40 | 22.5 | 0.016168 | 0.0041975 | 0.03259 | 0.00021519 | 0.0025925 | 0.011643 | 0.0044924 | 0.0081649 | 0.030893 | 0.0026922 | 0.03095 | 0.0059245 | 0.02252 | 0 | 0.023303 | 0.010809 | 0.015264 | 1.9805 | 3.5362 | 17.8438 | 57.3337 | 0.25073 | 2.0271 | 8.3452 | 0.80761 | 7.5047 | 0.14818
| T1-G | 41 | 23 | 0.014846 | 0.0056254 | 0.03141 | 0.00031172 | 0.0025837 | 0.011002 | 0.0047102 | 0.0075899 | 0.029979 | 0.0027207 | 0.030726 | 0.0062575 | 0.015771 | 0 | 0.031187 | 0.011192 | 0.016052 | 2.6195 | 3.4908 | 17.8034 | 56.9521 | 0.32099 | 2.0017 | 8.2401 | 0.80313 | 7.399 | 0.14744
| T1-G | 42 | 23.5 | 0.014566 | 0.0047871 | 0.031497 | 0.00041563 | 0.002993 | 0.011456 | 0.0045066 | 0.0077443 | 0.030307 | 0.0024059 | 0.031244 | 0.0056567 | 0.023742 | 0 | 0.026137 | 0.011714 | 0.014279 | 2.108 | 3.4837 | 17.9351 | 57.1674 | 0.31199 | 2.016 | 8.3221 | 0.80619 | 7.4776 | 0.14849
| T1-G | 43 | 24 | 0.011552 | 0.0062445 | 0.031194 | 0.00051507 | 0.002407 | 0.011017 | 0.0047064 | 0.0076342 | 0.029991 | 0.0027725 | 0.030351 | 0.0063075 | 0.013973 | 0 | 0.021949 | 0.01065 | 0.012768 | 2.3557 | 3.5315 | 17.8047 | 57.1647 | 0.26818 | 1.9949 | 8.2848 | 0.80831 | 7.4375 | 0.14568
Each single spot analysis was measured for 300 s using a beam energy of 50 kV and a current of 200 µA. | T1-G | 44 | 24.5 | 0.013889 | 0.0052556 | 0.03174 | 0.00053328 | 0.0018806 | 0.011121 | 0.0043699 | 0.0067141 | 0.029175 | 0.0027684 | 0.029837 | 0.0058763 | 0.030335 | 0 | 0.022497 | 0.01243 | 0.014438 | 2.5882 | 3.8126 | 17.8729 | 56.7485 | 0.24405 | 1.9915 | 8.223 | 0.79043 | 7.36 | 0.14596
| T1-G | 45 | 25 | 0.017525 | 0.005859 | 0.031997 | 0.00039027 | 0.0025424 | 0.01109 | 0.0042001 | 0.0080634 | 0.031207 | 0.0022282 | 0.030448 | 0.0059242 | 0.019401 | 0 | 0.031416 | 0.012428 | 0.014404 | 2.2068 | 3.4878 | 17.8063 | 57.2018 | 0.26566 | 2.0139 | 8.333 | 0.81095 | 7.4953 | 0.14939
| T1-G | 46 | 25.5 | 0.014686 | 0.005762 | 0.031352 | 0.00063543 | 0.0024008 | 0.011144 | 0.0048704 | 0.0084118 | 0.030688 | 0.001827 | 0.029613 | 0.006027 | 0.015655 | 0 | 0.026283 | 0.011329 | 0.01662 | 2.1239 | 3.3989 | 17.9207 | 57.3102 | 0.25388 | 2.0129 | 8.3215 | 0.80657 | 7.4882 | 0.14598
| T1-G | 47 | 26 | 0.014841 | 0.0054752 | 0.032396 | 0.00049572 | 0.0027452 | 0.011589 | 0.0047021 | 0.0069499 | 0.029741 | 0.0026624 | 0.029915 | 0.0050076 | 0.019651 | 0 | 0.026148 | 0.012415 | 0.013988 | 2.4829 | 3.4691 | 17.8148 | 57.0584 | 0.26895 | 2.0099 | 8.2819 | 0.8008 | 7.4469 | 0.1475
| T1-G | 48 | 26.5 | 0.013613 | 0.0058224 | 0.03259 | 0.00019451 | 0.0022051 | 0.011008 | 0.0048717 | 0.0077126 | 0.030699 | 0.0026068 | 0.031543 | 0.0060909 | 0.011151 | 0 | 0.027165 | 0.011582 | 0.015585 | 2.1335 | 3.4705 | 17.7868 | 57.3299 | 0.24389 | 2.016 | 8.3382 | 0.81883 | 7.4988 | 0.14892
| T1-G | 49 | 27 | 0.013672 | 0.0060677 | 0.031718 | 3.2439e-05 | 0.0025526 | 0.011158 | 0.0041365 | 0.0070983 | 0.030002 | 0.0025928 | 0.030863 | 0.0054881 | 0.016208 | 0 | 0.030455 | 0.012123 | 0.016598 | 2.3281 | 3.5565 | 17.7928 | 57.0856 | 0.29033 | 2.0029 | 8.3066 | 0.80709 | 7.462 | 0.14743
| T1-G | 50 | 27.5 | 0.015568 | 0.0061984 | 0.031881 | 0.00054246 | 0.0025047 | 0.010821 | 0.0048232 | 0.0077531 | 0.030547 | 0.0018023 | 0.03019 | 0.0055296 | 0.015339 | 0 | 0.027099 | 0.011332 | 0.014355 | 2.2681 | 3.5371 | 17.8064 | 57.1796 | 0.24301 | 2.0169 | 8.3046 | 0.80594 | 7.4739 | 0.14815
| T1-G | 51 | 28 | 0.015308 | 0.0053925 | 0.032115 | 0.00032072 | 0.0026974 | 0.011527 | 0.0047245 | 0.0077696 | 0.030646 | 0.0026474 | 0.03087 | 0.0056132 | 0.01719 | 0 | 0.026599 | 0.011973 | 0.013392 | 2.0759 | 3.4344 | 17.7919 | 57.3449 | 0.30588 | 2.0202 | 8.3395 | 0.81183 | 7.5072 | 0.14946
| T1-G | 52 | 28.5 | 0.017262 | 0.0047623 | 0.032202 | 0.00093337 | 0.0025183 | 0.010697 | 0.004326 | 0.0080578 | 0.030555 | 0.0016307 | 0.03092 | 0.006034 | 0.02412 | 0 | 0.027117 | 0.011435 | 0.014694 | 2.2776 | 3.4159 | 17.8428 | 57.162 | 0.25632 | 2.029 | 8.3391 | 0.80513 | 7.4962 | 0.1486
| T1-G | 53 | 29 | 0.016793 | 0.0049245 | 0.032614 | 0.00054936 | 0.0027618 | 0.011656 | 0.0039433 | 0.0075941 | 0.03054 | 0.0020601 | 0.030927 | 0.005862 | 0.020333 | 0 | 0.027309 | 0.011906 | 0.013554 | 2.2836 | 3.4526 | 17.824 | 57.1463 | 0.29133 | 2.0276 | 8.32 | 0.80515 | 7.4775 | 0.14859
| T1-G | 54 | 29.5 | 0.014849 | 0.0057712 | 0.031942 | 0.00071262 | 0.0022772 | 0.011072 | 0.004724 | 0.0079675 | 0.030938 | 0.0024401 | 0.030718 | 0.00569 | 0.019226 | 0 | 0.02815 | 0.011735 | 0.014236 | 2.1994 | 3.4359 | 17.8367 | 57.2133 | 0.30611 | 2.0358 | 8.3156 | 0.80385 | 7.4818 | 0.14907
| T1-G | 55 | 30 | 0.012658 | 0.0052858 | 0.032864 | 0.00027795 | 0.0021543 | 0.011334 | 0.0044306 | 0.0077392 | 0.031502 | 0.0017858 | 0.030779 | 0.0061214 | 0.011226 | 0 | 0.023921 | 0.01198 | 0.015832 | 2.1355 | 3.415 | 17.8022 | 57.3194 | 0.2752 | 2.025 | 8.3482 | 0.81702 | 7.5042 | 0.14846
| T1-G | 56 | 30.5 | 0.015262 | 0.0059622 | 0.032449 | 0.0002549 | 0.0029693 | 0.011886 | 0.0047249 | 0.0083902 | 0.030816 | 0.0026875 | 0.030462 | 0.0056628 | 0.021249 | 0 | 0.021906 | 0.011078 | 0.015769 | 2.2362 | 3.4558 | 17.8956 | 57.2227 | 0.26378 | 2.0149 | 8.2854 | 0.80086 | 7.4554 | 0.14793
| T1-G | 57 | 31 | 0.014512 | 0.0060477 | 0.032117 | 0.00064016 | 0.0024324 | 0.011122 | 0.0043238 | 0.0083996 | 0.030707 | 0.0017843 | 0.031934 | 0.0058795 | 0.020291 | 0 | 0.026131 | 0.011609 | 0.014814 | 2.1012 | 3.4825 | 17.8941 | 57.2198 | 0.28262 | 2.0099 | 8.3244 | 0.80786 | 7.5056 | 0.14925
| T1-G | 58 | 31.5 | 0.014909 | 0.0062857 | 0.032684 | 0 | 0.0024602 | 0.011093 | 0.0046945 | 0.0080941 | 0.030804 | 0.001968 | 0.031417 | 0.0061258 | 0.017388 | 0 | 0.024778 | 0.011626 | 0.015207 | 1.8093 | 3.3825 | 17.803 | 57.5356 | 0.28922 | 2.0207 | 8.3944 | 0.81767 | 7.5789 | 0.14914
| T1-G | 59 | 32 | 0.01575 | 0.0058797 | 0.031774 | 0.00043129 | 0.0023583 | 0.0118 | 0.0046549 | 0.008105 | 0.031202 | 0.0029296 | 0.031404 | 0.0052773 | 0.01998 | 0 | 0.026328 | 0.011152 | 0.013124 | 2.2336 | 3.3352 | 17.8849 | 57.2274 | 0.29266 | 2 | 8.3497 | 0.80713 | 7.4998 | 0.14744
| T1-G | 60 | 32.5 | 0.016003 | 0.0047323 | 0.031528 | 0.00023465 | 0.0025641 | 0.011603 | 0.0047898 | 0.0085012 | 0.031754 | 0.0029771 | 0.031013 | 0.0057434 | 0.015579 | 0 | 0.028451 | 0.011376 | 0.01397 | 2.1282 | 3.4462 | 17.9185 | 57.1965 | 0.2786 | 2.0196 | 8.3238 | 0.81445 | 7.5044 | 0.14896
| T1-G | 61 | 33 | 0.015273 | 0.0052489 | 0.03202 | 0.00035085 | 0.0026651 | 0.011265 | 0.0046498 | 0.0079938 | 0.030041 | 0.0020992 | 0.030963 | 0.0055465 | 0.014664 | 0 | 0.027848 | 0.011796 | 0.015635 | 2.3451 | 3.5191 | 17.8221 | 57.1185 | 0.28179 | 2.0081 | 8.2812 | 0.80509 | 7.4536 | 0.14733
| T1-G | 62 | 33.5 | 0.018266 | 0.0056633 | 0.032219 | 0.00062353 | 0.0027837 | 0.011968 | 0.004788 | 0.0080821 | 0.030819 | 0.0025836 | 0.032166 | 0.0062316 | 0.018397 | 0 | 0.028057 | 0.011188 | 0.013681 | 1.7895 | 3.3317 | 17.8597 | 57.5398 | 0.26619 | 2.0273 | 8.4162 | 0.81445 | 7.5758 | 0.1518
| T1-G | 63 | 34 | 0.015777 | 0.005211 | 0.03139 | 0.00026617 | 0.0024993 | 0.011282 | 0.0048899 | 0.0070638 | 0.030903 | 0.0033198 | 0.031962 | 0.0053278 | 0.017163 | 0 | 0.026084 | 0.011527 | 0.014003 | 2.2904 | 3.4666 | 17.8103 | 57.1898 | 0.29918 | 2.0002 | 8.2963 | 0.80846 | 7.4726 | 0.14751
| T1-G | 64 | 34.5 | 0.015321 | 0.0057923 | 0.031477 | 0.00056231 | 0.002311 | 0.011297 | 0.004495 | 0.0078462 | 0.029235 | 0.0022822 | 0.030732 | 0.0061712 | 0.0099881 | 0 | 0.031995 | 0.011657 | 0.014632 | 2.6319 | 3.5253 | 17.8228 | 56.9224 | 0.26518 | 1.9824 | 8.2577 | 0.80722 | 7.4222 | 0.14713
| T1-G | 65 | 35 | 0.015668 | 0.005586 | 0.031057 | 0.00057617 | 0.0024924 | 0.011422 | 0.004999 | 0.0077552 | 0.031021 | 0.0025371 | 0.030136 | 0.0057973 | 0.01821 | 0 | 0.029639 | 0.010633 | 0.014618 | 2.3085 | 3.4012 | 17.9033 | 57.1299 | 0.25988 | 2.01 | 8.3261 | 0.79832 | 7.4935 | 0.14724
| T1-G | 66 | 35.5 | 0.015363 | 0.005852 | 0.031244 | 0.0006677 | 0.002897 | 0.011512 | 0.0046591 | 0.0079427 | 0.031313 | 0.0024149 | 0.030656 | 0.0060462 | 0.017844 | 0 | 0.026319 | 0.011419 | 0.014248 | 2.2503 | 3.5088 | 17.8812 | 57.1157 | 0.26502 | 2.0097 | 8.3138 | 0.8088 | 7.4776 | 0.14867
| T1-G | 67 | 36 | 0.01531 | 0.0062483 | 0.03152 | 0.00023377 | 0.0024493 | 0.011468 | 0.0047842 | 0.0075753 | 0.030019 | 0.0013426 | 0.030612 | 0.0058418 | 0.0050203 | 0 | 0.031466 | 0.012165 | 0.013579 | 2.287 | 3.4192 | 17.7912 | 57.2386 | 0.2696 | 2.0189 | 8.3137 | 0.81112 | 7.4913 | 0.14965
| T1-G | 68 | 36.5 | 0.017005 | 0.0051124 | 0.032025 | 0.00032911 | 0.0025423 | 0.011039 | 0.0047042 | 0.0089373 | 0.031275 | 0.0018467 | 0.030779 | 0.0063889 | 0.019635 | 0 | 0.026218 | 0.010023 | 0.015252 | 2.2957 | 3.4366 | 17.7519 | 57.2322 | 0.26311 | 2.0189 | 8.3237 | 0.80734 | 7.4999 | 0.14761
| T1-G | 69 | 37 | 0.015229 | 0.0047408 | 0.032166 | 0.00025537 | 0.0026658 | 0.011589 | 0.0045088 | 0.0081145 | 0.031223 | 0.0024567 | 0.030988 | 0.0052003 | 0.019762 | 0 | 0.021371 | 0.011365 | 0.015634 | 2.1248 | 3.4125 | 17.7982 | 57.3166 | 0.30244 | 2.0201 | 8.3302 | 0.81128 | 7.5168 | 0.14981
| T1-G | 70 | 37.5 | 0.016461 | 0.0043733 | 0.031066 | 0.00061982 | 0.00268 | 0.011326 | 0.004594 | 0.0081049 | 0.030737 | 0.002097 | 0.03067 | 0.0051972 | 0.023904 | 0 | 0.023364 | 0.011979 | 0.013649 | 1.9 | 3.4013 | 17.8177 | 57.4534 | 0.28963 | 2.0402 | 8.3601 | 0.81049 | 7.5576 | 0.14868
| T1-G | 71 | 38 | 0.016587 | 0.0054629 | 0.031243 | 0.00041146 | 0.0022152 | 0.011158 | 0.0046846 | 0.0080736 | 0.030734 | 0.0027151 | 0.030922 | 0.0054754 | 0.0087331 | 0 | 0.033485 | 0.011385 | 0.015467 | 2.4335 | 3.5172 | 17.7912 | 57.0647 | 0.24711 | 1.9991 | 8.309 | 0.8137 | 7.4589 | 0.14673
| T1-G | 72 | 38.5 | 0.018558 | 0.004959 | 0.031559 | 0.00050893 | 0.0025649 | 0.010977 | 0.0045312 | 0.0076511 | 0.031385 | 0.0023198 | 0.030417 | 0.0056549 | 0.0096648 | 0 | 0.033971 | 0.011156 | 0.013545 | 2.1322 | 3.493 | 17.8912 | 57.202 | 0.2555 | 2.0122 | 8.3271 | 0.81331 | 7.5051 | 0.14899
| T1-G | 73 | 39 | 0.017754 | 0.0052765 | 0.031349 | 0.00022043 | 0.0021327 | 0.011217 | 0.0049722 | 0.0074262 | 0.03065 | 0.0024292 | 0.031061 | 0.0064771 | 0.015866 | 0 | 0.029781 | 0.010408 | 0.014921 | 2.2271 | 3.5816 | 17.8353 | 57.1682 | 0.25016 | 2.0038 | 8.2866 | 0.80069 | 7.4762 | 0.14846
| T1-G | 74 | 39.5 | 0.01557 | 0.0049998 | 0.030801 | 0.00057301 | 0.0023048 | 0.011167 | 0.0048723 | 0.0077474 | 0.030949 | 0.0026907 | 0.031841 | 0.0059156 | 0.021286 | 0 | 0.023618 | 0.010242 | 0.016005 | 2.4946 | 3.4968 | 17.7898 | 57.0543 | 0.248 | 2.0068 | 8.3 | 0.80075 | 7.4415 | 0.14682
| T1-G | 75 | 40 | 0.013851 | 0.0058495 | 0.031543 | 0.0002459 | 0.0026161 | 0.011521 | 0.0043282 | 0.0077509 | 0.030956 | 0.0029652 | 0.030593 | 0.0055167 | 0.016538 | 0 | 0.026933 | 0.011553 | 0.015611 | 2.5489 | 3.4102 | 17.8441 | 57.0372 | 0.26614 | 2.0085 | 8.2831 | 0.79996 | 7.4371 | 0.14632
| T1-G | 76 | 40.5 | 0.016749 | 0.0056318 | 0.031135 | 0.00063272 | 0.0027611 | 0.011538 | 0.0047563 | 0.0087295 | 0.031108 | 0.0025095 | 0.030985 | 0.0052556 | 0.018406 | 0 | 0.031041 | 0.01202 | 0.014043 | 1.9501 | 3.5536 | 17.8675 | 57.3331 | 0.24867 | 2.0202 | 8.3344 | 0.8072 | 7.5096 | 0.14836
| T1-G | 77 | 41 | 0.016087 | 0.006297 | 0.03192 | 0.00056095 | 0.002948 | 0.011855 | 0.0050393 | 0.0077185 | 0.031216 | 0.0018631 | 0.03064 | 0.0056286 | 0.019482 | 0 | 0.029961 | 0.010105 | 0.012912 | 1.6174 | 3.4601 | 17.8772 | 57.6031 | 0.26607 | 2.0308 | 8.3785 | 0.81445 | 7.5762 | 0.15184
| T1-G | 78 | 41.5 | 0.017261 | 0.0058541 | 0.031945 | 0.00086107 | 0.0024309 | 0.010835 | 0.0044658 | 0.0072195 | 0.029992 | 0.0021626 | 0.030856 | 0.0061417 | 0.024617 | 0 | 0.024766 | 0.010906 | 0.014269 | 2.1081 | 3.4773 | 17.8511 | 57.2598 | 0.25624 | 2.0135 | 8.3411 | 0.8056 | 7.5139 | 0.14878
| T1-G | 79 | 42 | 0.016377 | 0.006732 | 0.030472 | 0.00049909 | 0.0022092 | 0.011101 | 0.0045841 | 0.0074441 | 0.029999 | 0.0024468 | 0.031072 | 0.0055371 | 0.014266 | 0 | 0.031744 | 0.012118 | 0.014522 | 2.1964 | 3.5436 | 17.8135 | 57.2199 | 0.28119 | 2.0041 | 8.3124 | 0.81119 | 7.4477 | 0.14899
| T1-G | 80 | 42.5 | 0.017521 | 0.0049026 | 0.030377 | 0.00013137 | 0.0020775 | 0.011156 | 0.0042278 | 0.0082568 | 0.030979 | 0.0025805 | 0.029804 | 0.0054879 | 0.016215 | 0 | 0.02998 | 0.012114 | 0.015481 | 2.8632 | 3.568 | 17.7626 | 56.7783 | 0.32036 | 1.989 | 8.207 | 0.79928 | 7.3449 | 0.14613
| T1-G | 81 | 43 | 0.015031 | 0.0053441 | 0.032006 | 0.00028301 | 0.0025062 | 0.010855 | 0.0050669 | 0.007464 | 0.030087 | 0.0030912 | 0.031244 | 0.0055474 | 0.014903 | 0 | 0.029251 | 0.010578 | 0.015092 | 2.214 | 3.3898 | 17.875 | 57.254 | 0.26627 | 2.0247 | 8.3276 | 0.81056 | 7.4719 | 0.14785
| T1-G | 82 | 43.5 | 0.015972 | 0.0042807 | 0.031343 | 0.00039391 | 0.0019284 | 0.011609 | 0.004188 | 0.0083672 | 0.031858 | 0.0017981 | 0.030363 | 0.0062041 | 0.024957 | 0 | 0.021376 | 0.012826 | 0.01706 | 2.1658 | 3.4215 | 17.8804 | 57.2606 | 0.23549 | 2.0278 | 8.3334 | 0.80488 | 7.4971 | 0.14847
| T1-G | 83 | 44 | 0.017597 | 0.004997 | 0.031575 | 0.00052165 | 0.0027649 | 0.011918 | 0.004352 | 0.0077401 | 0.031496 | 0.002504 | 0.030755 | 0.005169 | 0.02658 | 0 | 0.024465 | 0.01217 | 0.013967 | 2.2409 | 3.369 | 17.7663 | 57.329 | 0.23271 | 2.0276 | 8.3498 | 0.80203 | 7.5059 | 0.14824
| T1-G | 84 | 44.5 | 0.016386 | 0.0050098 | 0.031309 | 0.00022283 | 0.0025359 | 0.010887 | 0.0042906 | 0.0075287 | 0.029625 | 0.0024463 | 0.030551 | 0.0058994 | 0.027838 | 0 | 0.024213 | 0.010736 | 0.015242 | 2.618 | 3.3363 | 17.8214 | 57.0698 | 0.26004 | 2.0104 | 8.2754 | 0.79615 | 7.4415 | 0.1462
| T1-G | 85 | 45 | 0.018444 | 0.0053639 | 0.031848 | 0.00040266 | 0.0027055 | 0.011763 | 0.0045061 | 0.0072816 | 0.030221 | 0.0028981 | 0.030501 | 0.0052793 | 0.01858 | 0 | 0.032201 | 0.012744 | 0.014175 | 2.1703 | 3.5691 | 17.8517 | 57.1729 | 0.26927 | 2.0144 | 8.308 | 0.80719 | 7.4605 | 0.14776
| T1-G | 86 | 45.5 | 0.013598 | 0.0066507 | 0.030601 | 0.00062961 | 0.0030559 | 0.011149 | 0.0048397 | 0.0080967 | 0.031066 | 0.0019388 | 0.029971 | 0.0055391 | 0.01413 | 0 | 0.028448 | 0.0095638 | 0.015678 | 2.2725 | 3.533 | 17.7414 | 57.2078 | 0.27618 | 2.0341 | 8.3031 | 0.80849 | 7.4614 | 0.14707
| T1-G | 87 | 46 | 0.016696 | 0.005844 | 0.031202 | 0.00022444 | 0.0023922 | 0.010952 | 0.0040164 | 0.0073099 | 0.030375 | 0.0024558 | 0.029665 | 0.0058526 | 0.020468 | 0 | 0.028224 | 0.011227 | 0.014367 | 2.1295 | 3.3918 | 17.8826 | 57.3132 | 0.2779 | 2.0271 | 8.3212 | 0.8014 | 7.4856 | 0.14848
| T1-G | 88 | 46.5 | 0.016041 | 0.0057976 | 0.031195 | 0.00040302 | 0.0026688 | 0.011706 | 0.0045805 | 0.0071668 | 0.030438 | 0.0021977 | 0.030137 | 0.0054506 | 0.011131 | 0 | 0.031583 | 0.011237 | 0.014561 | 2.1342 | 3.5003 | 17.8622 | 57.218 | 0.30246 | 2.0335 | 8.3037 | 0.81301 | 7.4668 | 0.14961
| T1-G | 89 | 47 | 0.016667 | 0.0052513 | 0.031838 | 0.00026329 | 0.0026408 | 0.01098 | 0.003784 | 0.0077556 | 0.031558 | 0.0018851 | 0.029971 | 0.0053343 | 0.020743 | 0 | 0.028303 | 0.011567 | 0.014385 | 2.0598 | 3.4701 | 17.8862 | 57.2548 | 0.26691 | 2.031 | 8.3446 | 0.80651 | 7.5079 | 0.14928
| T1-G | 90 | 47.5 | 0.014274 | 0.00713 | 0.030743 | 0.00059594 | 0.0028934 | 0.011367 | 0.0046859 | 0.007562 | 0.030722 | 0.002802 | 0.031419 | 0.0052892 | 0.022307 | 0 | 0.02566 | 0.011424 | 0.015702 | 2.4987 | 3.5199 | 17.9133 | 56.9724 | 0.25139 | 2.0071 | 8.2624 | 0.79624 | 7.4084 | 0.14563
| T1-G | 91 | 48 | 0.016561 | 0.0058214 | 0.031508 | 0.00089687 | 0.0025774 | 0.011324 | 0.0043825 | 0.0079936 | 0.031015 | 0.0018139 | 0.030299 | 0.0053388 | 0.023609 | 0 | 0.028822 | 0.011972 | 0.015029 | 2.2769 | 3.5189 | 17.8136 | 57.1755 | 0.25941 | 2.0229 | 8.2894 | 0.80464 | 7.4612 | 0.14852
| T1-G | 92 | 48.5 | 0.017136 | 0.0053955 | 0.031256 | 0.00046174 | 0.0023671 | 0.011313 | 0.0049834 | 0.007859 | 0.030235 | 0.0029823 | 0.030458 | 0.0061203 | 0.020117 | 0 | 0.027876 | 0.011127 | 0.014109 | 2.1317 | 3.4666 | 17.8597 | 57.3288 | 0.2332 | 2.028 | 8.3102 | 0.80339 | 7.4652 | 0.1493
| T1-G | 93 | 49 | 0.015428 | 0.0057567 | 0.030339 | 0 | 0.002775 | 0.010985 | 0.0043467 | 0.0075329 | 0.030292 | 0.0023113 | 0.030138 | 0.0061086 | 0.023432 | 0 | 0.028202 | 0.010907 | 0.013603 | 2.4537 | 3.5031 | 17.7632 | 57.1159 | 0.27952 | 2.0091 | 8.2728 | 0.80212 | 7.4306 | 0.14769
| T1-G | 94 | 49.5 | 0.017654 | 0.0062531 | 0.031894 | 0.00057004 | 0.0024588 | 0.011321 | 0.004459 | 0.0080298 | 0.030918 | 0.0017511 | 0.029814 | 0.0052121 | 0.020215 | 0 | 0.02899 | 0.01087 | 0.013712 | 2.1224 | 3.4151 | 17.9068 | 57.3129 | 0.24945 | 2.0252 | 8.311 | 0.80159 | 7.4818 | 0.14968
| T1-G | 95 | 50 | 0.014235 | 0.0058275 | 0.031517 | 0 | 0.0025211 | 0.011075 | 0.0044762 | 0.0072045 | 0.030771 | 0.0029692 | 0.031061 | 0.0058409 | 0.014573 | 0 | 0.02703 | 0.010287 | 0.016751 | 2.2032 | 3.4698 | 17.9116 | 57.1858 | 0.29271 | 2.0043 | 8.3112 | 0.80869 | 7.4492 | 0.14744
| T1-G | 96 | 50.5 | 0.01742 | 0.0053305 | 0.031782 | 0.00076006 | 0.0026502 | 0.01154 | 0.0043994 | 0.0073093 | 0.029869 | 0.0022673 | 0.030224 | 0.0056311 | 0.015989 | 0 | 0.029847 | 0.011244 | 0.014745 | 2.4332 | 3.4886 | 17.8813 | 57.0718 | 0.25282 | 2.0107 | 8.2783 | 0.80007 | 7.4147 | 0.14754
| T1-G | 97 | 51 | 0.016371 | 0.0057882 | 0.030257 | 0.00072673 | 0.0021038 | 0.011012 | 0.0044269 | 0.0084477 | 0.030923 | 0.0022391 | 0.030151 | 0.0056066 | 0.024409 | 0 | 0.027183 | 0.011877 | 0.014916 | 2.3541 | 3.553 | 17.8746 | 57.1037 | 0.25177 | 2.0057 | 8.2696 | 0.79992 | 7.4147 | 0.14643
| T1-G | 98 | 51.5 | 0.016929 | 0.0056209 | 0.031299 | 0.00028303 | 0.0026106 | 0.01098 | 0.0042033 | 0.0078039 | 0.030922 | 0.0027029 | 0.030225 | 0.0056077 | 0.0138 | 0 | 0.032034 | 0.012382 | 0.013973 | 2.184 | 3.5034 | 17.8672 | 57.2003 | 0.28862 | 2.0083 | 8.3102 | 0.81288 | 7.4561 | 0.1477
| T1-G | 99 | 52 | 0.020595 | 0.0054427 | 0.031141 | 0 | 0.0028681 | 0.011788 | 0.0049506 | 0.0081983 | 0.029953 | 0.0023682 | 0.02995 | 0.0062248 | 0.010765 | 0 | 0.035753 | 0.0117 | 0.015777 | 2.1705 | 3.6168 | 17.8546 | 57.1446 | 0.28724 | 2.0172 | 8.2798 | 0.8057 | 7.447 | 0.14908
| T1-G | 100 | 52.5 | 0.017434 | 0.0058985 | 0.03124 | 0.00058943 | 0.0023871 | 0.010504 | 0.0040502 | 0.0079555 | 0.030687 | 0.0021219 | 0.030482 | 0.0053574 | 0.01767 | 0 | 0.032044 | 0.011391 | 0.014114 | 2.3428 | 3.4612 | 17.8027 | 57.2307 | 0.28011 | 2.0121 | 8.2639 | 0.80439 | 7.4295 | 0.14854
| T1-G | 101 | 53 | 0.017516 | 0.0049678 | 0.031666 | 0.00041782 | 0.0022259 | 0.011465 | 0.0045551 | 0.0078857 | 0.030431 | 0.0019961 | 0.03154 | 0.0049099 | 0.019669 | 0 | 0.027781 | 0.012155 | 0.015548 | 2.421 | 3.4677 | 17.8115 | 57.1777 | 0.26869 | 2.0076 | 8.2647 | 0.7944 | 7.4149 | 0.147
| T1-G | 102 | 53.5 | 0.016402 | 0.0053723 | 0.031328 | 0.00047507 | 0.0028028 | 0.010767 | 0.0048314 | 0.007662 | 0.030268 | 0.0021621 | 0.030473 | 0.0057885 | 0.019734 | 0 | 0.029339 | 0.012365 | 0.013946 | 2.2095 | 3.525 | 17.8058 | 57.2319 | 0.30859 | 2.0103 | 8.2835 | 0.8043 | 7.45 | 0.1474
| T1-G | 103 | 54 | 0.012968 | 0.0056513 | 0.030412 | 0.00033113 | 0.0025911 | 0.011105 | 0.0048355 | 0.0070783 | 0.030437 | 0.0017679 | 0.030069 | 0.0053537 | 0.017147 | 0 | 0.023685 | 0.011525 | 0.01455 | 2.6498 | 3.5263 | 17.8819 | 56.916 | 0.27545 | 1.9914 | 8.2318 | 0.79791 | 7.3751 | 0.14476
| T1-G | 104 | 54.5 | 0.01752 | 0.0048998 | 0.030555 | 0.00058212 | 0.0023473 | 0.011136 | 0.0050296 | 0.0075682 | 0.030512 | 0.0029205 | 0.03082 | 0.00585 | 0.014808 | 0 | 0.02957 | 0.010653 | 0.013852 | 2.3614 | 3.3921 | 17.9238 | 57.1422 | 0.2698 | 2.0101 | 8.2852 | 0.80251 | 7.4457 | 0.14871
| T1-G | 105 | 55 | 0.017367 | 0.0054908 | 0.030526 | 0 | 0.0022942 | 0.011 | 0.0043153 | 0.0079238 | 0.031374 | 0.0028666 | 0.029145 | 0.0057848 | 0.01928 | 0 | 0.027221 | 0.0099695 | 0.014298 | 2.4013 | 3.454 | 17.9343 | 57.094 | 0.26129 | 2.0084 | 8.2608 | 0.80074 | 7.4179 | 0.14841
| T1-G | 106 | 55.5 | 0.016485 | 0.0069899 | 0.030973 | 0.00055975 | 0.0027886 | 0.011141 | 0.0040125 | 0.0073275 | 0.031832 | 0.0031464 | 0.03171 | 0.0056445 | 0.013875 | 0 | 0.029616 | 0.011579 | 0.015061 | 1.9021 | 3.4507 | 17.8778 | 57.3978 | 0.25942 | 2.0249 | 8.3776 | 0.81452 | 7.5236 | 0.14861
| T1-G | 107 | 56 | 0.016538 | 0.0070035 | 0.031001 | 0.00017571 | 0.0028891 | 0.011219 | 0.0046729 | 0.0077165 | 0.030819 | 0.0018107 | 0.031111 | 0.0055293 | 0.011152 | 0 | 0.032085 | 0.010608 | 0.012879 | 2.4405 | 3.5141 | 17.8106 | 57.159 | 0.26073 | 1.9986 | 8.2416 | 0.79884 | 7.4111 | 0.1477
| T1-G | 108 | 56.5 | 0.015591 | 0.0057177 | 0.030953 | 0.000324 | 0.0027489 | 0.011476 | 0.0042708 | 0.0084394 | 0.031432 | 0.0025204 | 0.031149 | 0.0053693 | 0.015601 | 0 | 0.03119 | 0.011522 | 0.015272 | 2.0206 | 3.3241 | 17.9466 | 57.395 | 0.24468 | 2.0077 | 8.3541 | 0.81755 | 7.5171 | 0.14894
| T1-G | 109 | 57 | 0.016459 | 0.0050677 | 0.030586 | 0.00014274 | 0.0026357 | 0.011463 | 0.0041864 | 0.0074722 | 0.03016 | 0.002326 | 0.030346 | 0.0051041 | 0.019186 | 0 | 0.024845 | 0.011839 | 0.015699 | 2.5052 | 3.5459 | 17.8527 | 56.9955 | 0.3277 | 1.99 | 8.2275 | 0.79389 | 7.3975 | 0.14651
| T1-G | 110 | 57.5 | 0.014943 | 0.0053102 | 0.03141 | 8.239e-05 | 0.0021908 | 0.011646 | 0.0042154 | 0.0067663 | 0.03049 | 0.0023714 | 0.030816 | 0.0058733 | 0.011052 | 0 | 0.032364 | 0.012914 | 0.014898 | 2.0774 | 3.4699 | 17.8968 | 57.3015 | 0.29323 | 2.0133 | 8.3107 | 0.80884 | 7.462 | 0.14888
| T1-G | 111 | 58 | 0.017113 | 0.005671 | 0.031421 | 4.6915e-05 | 0.0024751 | 0.010512 | 0.0042416 | 0.0080646 | 0.031197 | 0.0026182 | 0.030631 | 0.0060693 | 0.018209 | 0 | 0.027713 | 0.011873 | 0.015642 | 2.0836 | 3.5585 | 17.7964 | 57.356 | 0.24683 | 2.0169 | 8.2937 | 0.81084 | 7.4632 | 0.1506
| T1-G | 112 | 58.5 | 0.016428 | 0.0057035 | 0.031474 | 0.00053904 | 0.0021493 | 0.011114 | 0.0045664 | 0.007547 | 0.031184 | 0.0026707 | 0.031598 | 0.0051662 | 0.0086991 | 0 | 0.033591 | 0.010236 | 0.015564 | 2.0622 | 3.419 | 17.9138 | 57.3305 | 0.28936 | 2.0137 | 8.3177 | 0.81446 | 7.4721 | 0.14874
| T1-G | 113 | 59 | 0.016281 | 0.0063087 | 0.030878 | 0.0004213 | 0.0029005 | 0.011148 | 0.0038394 | 0.0070075 | 0.030388 | 0.0025661 | 0.031084 | 0.0052868 | 0.008559 | 0 | 0.031822 | 0.011765 | 0.013815 | 2.1288 | 3.5667 | 17.8897 | 57.2552 | 0.27561 | 2.0087 | 8.2692 | 0.81033 | 7.4343 | 0.14735
| T1-G | 114 | 59.5 | 0.015156 | 0.0064833 | 0.031016 | 0.0001279 | 0.0024861 | 0.010928 | 0.0051081 | 0.0082465 | 0.031046 | 0.0018997 | 0.029994 | 0.0055405 | 0.0090978 | 0 | 0.029502 | 0.0096106 | 0.015841 | 2.4022 | 3.5193 | 17.8305 | 57.1477 | 0.21349 | 2.0232 | 8.2665 | 0.80635 | 7.4315 | 0.14709
| T1-G | 115 | 60 | 0.01692 | 0.0056703 | 0.032052 | 0.00077871 | 0.0022109 | 0.011488 | 0.0044471 | 0.0074746 | 0.030942 | 0.0024194 | 0.030695 | 0.0056547 | 0.015639 | 0 | 0.031832 | 0.012158 | 0.014299 | 2.1438 | 3.5154 | 17.9172 | 57.2651 | 0.25102 | 2.0017 | 8.3017 | 0.80204 | 7.4297 | 0.14766
| T1-G | 116 | 60.5 | 0.017352 | 0.0060318 | 0.030496 | 0.00035091 | 0.0026925 | 0.010652 | 0.0043708 | 0.0070911 | 0.030402 | 0.0024316 | 0.031029 | 0.0049668 | 0.0091576 | 0 | 0.030308 | 0.010774 | 0.014887 | 2.3885 | 3.5321 | 17.8151 | 57.1479 | 0.29856 | 1.9967 | 8.2534 | 0.80552 | 7.4009 | 0.14813
| T1-G | 117 | 61 | 0.019741 | 0.005226 | 0.029951 | 0.00040441 | 0.0025089 | 0.011132 | 0.0045703 | 0.0076141 | 0.030344 | 0.0031213 | 0.03152 | 0.0049994 | 0.022108 | 0 | 0.029141 | 0.010636 | 0.013588 | 2.3044 | 3.6088 | 17.8083 | 57.1476 | 0.28421 | 1.992 | 8.2699 | 0.79509 | 7.4147 | 0.14837
| T1-G | 118 | 61.5 | 0.012892 | 0.0071296 | 0.03103 | 0.00053122 | 0.0026615 | 0.01098 | 0.004492 | 0.0080206 | 0.029466 | 0.001575 | 0.030977 | 0.0060342 | 0.011665 | 0 | 0.030175 | 0.010942 | 0.013769 | 2.2964 | 3.5415 | 17.7988 | 57.2468 | 0.24254 | 2.0057 | 8.2748 | 0.80737 | 7.4273 | 0.14637
| T1-G | 119 | 62 | 0.01768 | 0.0059084 | 0.031187 | 0.00054135 | 0.0029485 | 0.011375 | 0.0042548 | 0.0075879 | 0.030001 | 0.0024468 | 0.030353 | 0.0057088 | 0.015053 | 0 | 0.032562 | 0.011608 | 0.014523 | 2.1076 | 3.5388 | 17.8558 | 57.3142 | 0.26851 | 1.9921 | 8.2984 | 0.80631 | 7.446 | 0.14863
| T1-G | 120 | 62.5 | 0.015997 | 0.0059959 | 0.030153 | 0.00095409 | 0.0024129 | 0.011161 | 0.0051194 | 0.0075836 | 0.029721 | 0.0018093 | 0.02953 | 0.0056539 | 0.022112 | 0 | 0.026867 | 0.011422 | 0.014423 | 2.1773 | 3.5284 | 17.8567 | 57.2895 | 0.28857 | 2.0067 | 8.2581 | 0.80363 | 7.4227 | 0.14759
| T1-G | 121 | 63 | 0.019261 | 0.0053186 | 0.030937 | 0.00052739 | 0.0027519 | 0.011069 | 0.0051491 | 0.0072592 | 0.031028 | 0.0023043 | 0.030145 | 0.0055582 | 0.019022 | 0 | 0.02773 | 0.010365 | 0.013961 | 2.1817 | 3.5204 | 17.8921 | 57.2892 | 0.23218 | 2.0154 | 8.2693 | 0.80199 | 7.4277 | 0.14769
| T1-G | 122 | 63.5 | 0.017557 | 0.0054578 | 0.030463 | 0.00036678 | 0.0022117 | 0.011155 | 0.0043149 | 0.0078059 | 0.031333 | 0.0018311 | 0.030412 | 0.0057604 | 0.024949 | 0 | 0.02807 | 0.012355 | 0.014023 | 2.1609 | 3.3776 | 17.9022 | 57.299 | 0.29687 | 2.0242 | 8.2966 | 0.80054 | 7.4643 | 0.14984
| T1-G | 123 | 64 | 0.017471 | 0.0058799 | 0.031668 | 0.00070378 | 0.0030574 | 0.011489 | 0.004888 | 0.0088122 | 0.030743 | 0.0017634 | 0.030368 | 0.0058183 | 0 | 0 | 0.037442 | 0.011354 | 0.01428 | 2.4097 | 3.4034 | 17.8623 | 57.2148 | 0.26468 | 1.9961 | 8.2689 | 0.80843 | 7.4092 | 0.14669
| Sample | Analysis # | t after start (h) | V | Cr | Co | Ni | Cu | Zn | Ga | Rb | Sr | Y | Zr | Nb | Ba | La | Ce | Ta | Th | Na2O | MgO | Al2O3 | SiO2 | P2O5 | K2O | CaO | TiO2 | Fe2O3 | MnO
| K5 | 1 | 0 | 0 | 0.00053185 | 0.0099127 | 0.00044909 | 0.0015032 | 0.017637 | 0.0055685 | 0.0097037 | 0.0059538 | 0.0087534 | 0.052487 | 0.017533 | 0.054913 | 0.0025911 | 0.018989 | 0.0075932 | 0.01171 | 3.8728 | 0.19966 | 14.1486 | 73.0391 | 0.078918 | 4.8037 | 0.99081 | 0.19392 | 2.3729 | 0.073799
| K5 | 2 | 0.5 | 0 | 0.00083733 | 0.010219 | 0.00030411 | 0.001447 | 0.017089 | 0.0058841 | 0.0098755 | 0.0064102 | 0.0086178 | 0.05286 | 0.017761 | 0.057007 | 0.001979 | 0.017799 | 0.0075838 | 0.011032 | 3.3935 | 0.15926 | 14.1761 | 73.4287 | 0.084034 | 4.8586 | 1.0093 | 0.1952 | 2.3943 | 0.074229
| K5 | 3 | 1 | 0 | 0.00093338 | 0.010142 | 0.00045943 | 0.0014931 | 0.017452 | 0.0061961 | 0.0094727 | 0.0058252 | 0.0082231 | 0.051966 | 0.016698 | 0.053061 | 0.0037431 | 0.018966 | 0.0076883 | 0.01087 | 3.9309 | 0.2294 | 14.1121 | 73.0374 | 0.044654 | 4.8047 | 0.98709 | 0.19515 | 2.3616 | 0.073827
| K5 | 4 | 1.5 | 0 | 0.00059011 | 0.010618 | 9.275e-05 | 0.0015415 | 0.017714 | 0.0056549 | 0.0096131 | 0.0054334 | 0.0080433 | 0.052651 | 0.018352 | 0.055974 | 0.0041805 | 0.018334 | 0.0079901 | 0.011631 | 3.4404 | 0.17912 | 14.1175 | 73.4321 | 0.094737 | 4.8442 | 1.0037 | 0.19509 | 2.3913 | 0.073591
| K5 | 5 | 2 | 0 | 0.00098522 | 0.0099953 | 0.00049483 | 0.0018727 | 0.017672 | 0.0055959 | 0.010044 | 0.0061256 | 0.008278 | 0.051875 | 0.017362 | 0.058183 | 0.0012845 | 0.019527 | 0.0081239 | 0.011318 | 3.6717 | 0.19159 | 14.1521 | 73.2106 | 0.090601 | 4.822 | 0.99213 | 0.19215 | 2.3752 | 0.073337
| K5 | 6 | 2.5 | 0 | 0.00036363 | 0.010165 | 0.00056152 | 0.0012207 | 0.017569 | 0.0053858 | 0.0088777 | 0.0060748 | 0.0085058 | 0.053689 | 0.018395 | 0.057609 | 0.0037896 | 0.017752 | 0.0079377 | 0.011092 | 3.8365 | 0.15308 | 14.1203 | 73.1061 | 0.084597 | 4.8431 | 0.99106 | 0.19184 | 2.3713 | 0.073111
| K5 | 7 | 3 | 0 | 0.00046285 | 0.010633 | 0.0007056 | 0.0017595 | 0.017593 | 0.0059143 | 0.0097232 | 0.0058263 | 0.0088692 | 0.053502 | 0.016978 | 0.056239 | 0.0028561 | 0.019834 | 0.0079615 | 0.011133 | 3.4426 | 0.15708 | 14.2678 | 73.3476 | 0.059831 | 4.8506 | 0.99326 | 0.19616 | 2.3809 | 0.074185
| K5 | 8 | 3.5 | 0 | 0.0015824 | 0.010087 | 0.00042429 | 0.0011674 | 0.017075 | 0.0055316 | 0.0096495 | 0.0059331 | 0.0091142 | 0.053859 | 0.018136 | 0.052468 | 0.0015266 | 0.017602 | 0.0074939 | 0.010776 | 3.5274 | 0.15075 | 14.1245 | 73.3894 | 0.063159 | 4.8607 | 0.99876 | 0.19949 | 2.3895 | 0.07381
| K5 | 9 | 4 | 0 | 0.00078022 | 0.010949 | 0.00054144 | 0.0016201 | 0.017328 | 0.0056403 | 0.0090083 | 0.0057268 | 0.008234 | 0.052445 | 0.017559 | 0.059384 | 0.0022598 | 0.018808 | 0.0080493 | 0.011526 | 3.7067 | 0.22816 | 14.1675 | 73.1324 | 0.071871 | 4.8421 | 0.98874 | 0.19301 | 2.3656 | 0.074041
| K5 | 10 | 4.5 | 0 | 0.00052018 | 0.009879 | 0.00053114 | 0.0017552 | 0.017611 | 0.0056312 | 0.0091893 | 0.0058365 | 0.0087887 | 0.052533 | 0.017194 | 0.058784 | 0.0029327 | 0.021099 | 0.0063988 | 0.010462 | 3.9801 | 0.2606 | 14.0407 | 73.017 | 0.062292 | 4.8058 | 0.98257 | 0.19251 | 2.3543 | 0.07491
| K5 | 11 | 5 | 0 | 0.00084898 | 0.010682 | 0.00091438 | 0.0011319 | 0.017378 | 0.0058457 | 0.0088862 | 0.0057041 | 0.0082791 | 0.052793 | 0.017556 | 0.053346 | 0.0041871 | 0.022406 | 0.0082114 | 0.010676 | 3.804 | 0.1784 | 14.156 | 73.0901 | 0.051915 | 4.8465 | 0.99723 | 0.19732 | 2.3761 | 0.073595
| K5 | 12 | 5.5 | 0 | 0.00055868 | 0.010739 | 0.00053836 | 0.0014206 | 0.01753 | 0.0059019 | 0.0091946 | 0.0054788 | 0.0078973 | 0.052238 | 0.017535 | 0.059055 | 0.0034839 | 0.015663 | 0.0091813 | 0.010073 | 3.7375 | 0.17922 | 14.1223 | 73.2118 | 0.079479 | 4.8112 | 0.99286 | 0.19266 | 2.3719 | 0.074589
| K5 | 13 | 6 | 0 | 0.00011562 | 0.010497 | 0.0007124 | 0.0016309 | 0.017473 | 0.0058014 | 0.0087932 | 0.0055806 | 0.0083907 | 0.05314 | 0.017349 | 0.058494 | 0.0031407 | 0.014518 | 0.0078003 | 0.010486 | 3.6861 | 0.17282 | 14.1226 | 73.2912 | 0.054795 | 4.8225 | 0.99219 | 0.19185 | 2.3684 | 0.073644
| K5 | 14 | 6.5 | 0 | 0.00025303 | 0.010303 | 0.0004591 | 0.0011179 | 0.017472 | 0.0059094 | 0.0099889 | 0.005939 | 0.0085496 | 0.05214 | 0.018009 | 0.0592 | 0.0041255 | 0.019748 | 0.0077189 | 0.011037 | 3.6934 | 0.18002 | 14.1638 | 73.174 | 0.081788 | 4.8481 | 0.98369 | 0.19603 | 2.3724 | 0.074806
| K5 | 15 | 17.55 | 0 | 0.0011652 | 0.010431 | 0.00061568 | 0.0015431 | 0.01734 | 0.0056832 | 0.0094522 | 0.0057754 | 0.0084559 | 0.052708 | 0.018551 | 0.055477 | 0.0028673 | 0.020512 | 0.0070181 | 0.0096727 | 3.804 | 0.19081 | 14.1336 | 73.1125 | 0.069408 | 4.8326 | 0.99874 | 0.19376 | 2.3636 | 0.073628
| K5 | 16 | 18.05 | 0 | 0.0011 | 0.010486 | 0.00089104 | 0.0015979 | 0.017537 | 0.0056 | 0.0096533 | 0.0062426 | 0.0086972 | 0.052291 | 0.01723 | 0.052762 | 0.0011832 | 0.017524 | 0.009347 | 0.011237 | 3.8468 | 0.2156 | 14.1721 | 72.9952 | 0.11429 | 4.8155 | 0.98392 | 0.19679 | 2.3632 | 0.073216
| K5 | 17 | 18.8667 | 0 | 0.00098681 | 0.010497 | 0.00050667 | 0.0018977 | 0.017385 | 0.0056588 | 0.0091937 | 0.005916 | 0.0086267 | 0.053348 | 0.017741 | 0.060385 | 0.0034957 | 0.022553 | 0.0070817 | 0.011103 | 3.9601 | 0.17505 | 14.0435 | 73.0714 | 0.078769 | 4.8157 | 0.98783 | 0.1899 | 2.3666 | 0.074753
| K5 | 18 | 24 | 0 | 0.00049982 | 0.010552 | 0.00047115 | 0.0017191 | 0.017395 | 0.0061061 | 0.0096119 | 0.0057648 | 0.0086039 | 0.053959 | 0.018393 | 0.060606 | 0.0045096 | 0.01676 | 0.0067195 | 0.01022 | 3.841 | 0.28591 | 14.0829 | 73.0236 | 0.059577 | 4.8351 | 0.99283 | 0.19042 | 2.3824 | 0.0745
| K5 | 19 | 24.5 | 0 | 0 | 0.011207 | 0.00035808 | 0.0014159 | 0.01766 | 0.0056252 | 0.0097745 | 0.0059946 | 0.0089235 | 0.053268 | 0.01728 | 0.055869 | 0.0014897 | 0.015865 | 0.0079517 | 0.011143 | 3.9737 | 0.31334 | 14.0687 | 72.9491 | 0.059875 | 4.8031 | 0.98214 | 0.1927 | 2.3597 | 0.073833
| K5 | 20 | 25 | 0 | 0.0001835 | 0.011118 | 0.00059751 | 0.0013642 | 0.017606 | 0.0060386 | 0.0092797 | 0.0056469 | 0.0085669 | 0.054336 | 0.017451 | 0.061059 | 0.0059675 | 0.016329 | 0.0074725 | 0.011795 | 3.4659 | 0.19517 | 14.0516 | 73.4451 | 0.057536 | 4.8757 | 1.0028 | 0.18955 | 2.4071 | 0.074774
| K5 | 21 | 25.5 | 0 | 0.0011314 | 0.009728 | 0.0011316 | 0.0020569 | 0.017999 | 0.0059313 | 0.0097744 | 0.0059401 | 0.0087696 | 0.052723 | 0.01882 | 0.054016 | 0.0036998 | 0.020492 | 0.0080439 | 0.011305 | 3.6992 | 0.20105 | 14.0854 | 73.2094 | 0.070782 | 4.8356 | 1.007 | 0.20174 | 2.3843 | 0.074034
| K5 | 22 | 26 | 0 | 0.00087924 | 0.010129 | 0.00050735 | 0.0018164 | 0.017441 | 0.0053257 | 0.0092197 | 0.0058072 | 0.0086109 | 0.052847 | 0.017884 | 0.053425 | 0.0029172 | 0.021233 | 0.0072298 | 0.010342 | 3.8124 | 0.15953 | 14.1648 | 73.0719 | 0.083439 | 4.8633 | 0.99032 | 0.19618 | 2.3574 | 0.075162
| K5 | 23 | 26.5 | 0 | 0.0010748 | 0.0099021 | 0.00040499 | 0.0012756 | 0.01744 | 0.0057533 | 0.0093665 | 0.0055588 | 0.0086091 | 0.052828 | 0.018399 | 0.052014 | 0.00080006 | 0.019747 | 0.0073889 | 0.011082 | 3.5958 | 0.20205 | 14.1917 | 73.2086 | 0.076804 | 4.8527 | 0.99945 | 0.19907 | 2.3788 | 0.07341
| K5 | 24 | 27 | 0 | 0.00072134 | 0.010246 | 0.00039431 | 0.00163 | 0.017554 | 0.0060372 | 0.0094396 | 0.0057745 | 0.0091201 | 0.053074 | 0.017725 | 0.050095 | 0.0028036 | 0.026021 | 0.0070211 | 0.011268 | 3.746 | 0.15074 | 14.1129 | 73.1556 | 0.10699 | 4.8583 | 0.99105 | 0.1996 | 2.3746 | 0.075168
| K5 | 25 | 27.5 | 0 | 0.0010067 | 0.010019 | 0.00079748 | 0.0017738 | 0.017463 | 0.0059145 | 0.0089425 | 0.0055289 | 0.0087799 | 0.052929 | 0.018244 | 0.05359 | 0.0033582 | 0.024433 | 0.0070581 | 0.011185 | 3.6054 | 0.29928 | 14.2199 | 73.1109 | 0.07045 | 4.8357 | 0.99632 | 0.19702 | 2.3595 | 0.074509
| K5 | 26 | 28 | 0 | 0.00077589 | 0.0098757 | 0.00072279 | 0.0018454 | 0.017593 | 0.005511 | 0.0089602 | 0.0053478 | 0.0086941 | 0.052405 | 0.018158 | 0.050517 | 0.0019188 | 0.018872 | 0.0077058 | 0.010608 | 3.5996 | 0.20276 | 14.0308 | 73.3737 | 0.048929 | 4.8818 | 0.99441 | 0.19878 | 2.3768 | 0.072871
| K5 | 27 | 28.5 | 0 | 0.0010203 | 0.010127 | 0.00019301 | 0.0015985 | 0.017361 | 0.0058852 | 0.0094913 | 0.0065159 | 0.0092273 | 0.053018 | 0.017214 | 0.052143 | 0.0020764 | 0.021891 | 0.008304 | 0.010032 | 3.8928 | 0.12796 | 14.1468 | 73.0678 | 0.081376 | 4.8398 | 0.98663 | 0.19699 | 2.3603 | 0.073406
| K5 | 28 | 29 | 0 | 0.0013714 | 0.010233 | 0.00076082 | 0.0015812 | 0.017195 | 0.005825 | 0.0098432 | 0.0062442 | 0.0083836 | 0.052964 | 0.018255 | 0.047283 | 0.0037111 | 0.022382 | 0.0069143 | 0.011777 | 3.8033 | 0.20011 | 14.0824 | 73.164 | 0.061146 | 4.8382 | 0.98762 | 0.20232 | 2.3628 | 0.073377
| K5 | 29 | 29.5 | 0 | 0.00082304 | 0.01012 | 0.00055916 | 0.0014431 | 0.017299 | 0.0055229 | 0.0089929 | 0.0055423 | 0.0091998 | 0.053539 | 0.018233 | 0.055157 | 0.0028599 | 0.026625 | 0.008709 | 0.012571 | 3.4865 | 0.10542 | 14.0996 | 73.4584 | 0.05534 | 4.8863 | 1.0057 | 0.20193 | 2.3887 | 0.07501
| K5 | 30 | 30 | 0 | 0.0013538 | 0.010046 | 0.00059101 | 0.0016022 | 0.017503 | 0.0057962 | 0.0095607 | 0.0056537 | 0.0086669 | 0.054184 | 0.017417 | 0.050399 | 0.0024043 | 0.021196 | 0.0083601 | 0.011218 | 3.5437 | 0.16452 | 14.137 | 73.3163 | 0.06991 | 4.8865 | 0.99515 | 0.20285 | 2.3831 | 0.075056
| K5 | 31 | 30.5 | 0 | 0.00077468 | 0.0099522 | 0.00051747 | 0.00094469 | 0.017664 | 0.0055596 | 0.0089869 | 0.0054958 | 0.0085533 | 0.053196 | 0.01855 | 0.053456 | 0.0028353 | 0.022834 | 0.0085383 | 0.0098187 | 3.6561 | 0.1702 | 14.0842 | 73.2631 | 0.070632 | 4.8658 | 1.0023 | 0.20136 | 2.3833 | 0.075473
| K5 | 32 | 31 | 0 | 0.0011809 | 0.010015 | 0.00046788 | 0.0018144 | 0.017734 | 0.0057058 | 0.0092595 | 0.0057751 | 0.0080769 | 0.053089 | 0.017766 | 0.049399 | 0.0022628 | 0.02168 | 0.0068545 | 0.010497 | 3.9689 | 0.19883 | 14.0702 | 73.0253 | 0.05332 | 4.825 | 0.99756 | 0.19987 | 2.3658 | 0.073671
| K5 | 33 | 31.5 | 0 | 0.0015613 | 0.01024 | 0.00056036 | 0.0014656 | 0.017111 | 0.0052238 | 0.0090859 | 0.0055538 | 0.0083205 | 0.052384 | 0.017652 | 0.052543 | 0.0017174 | 0.020938 | 0.0083702 | 0.010312 | 3.9553 | 0.20436 | 14.1525 | 72.9234 | 0.083849 | 4.8401 | 0.98743 | 0.19698 | 2.3591 | 0.073794
| K5 | 34 | 32 | 0 | 0.0011432 | 0.0099849 | 0.00084687 | 0.0016293 | 0.017371 | 0.0059027 | 0.0097844 | 0.0055865 | 0.0088103 | 0.052872 | 0.018524 | 0.053274 | 0.0030507 | 0.021801 | 0.0077149 | 0.011864 | 3.5436 | 0.20675 | 14.0991 | 73.3053 | 0.072169 | 4.8838 | 0.99878 | 0.19989 | 2.3859 | 0.074561
| K5 | 35 | 44.2333 | 0 | 0.0010196 | 0.0094741 | 0.00026672 | 0.0014528 | 0.017581 | 0.005865 | 0.0096963 | 0.0061881 | 0.0086445 | 0.054607 | 0.018646 | 0.052895 | 0.0014901 | 0.019496 | 0.0068082 | 0.011615 | 3.5098 | 0.18684 | 14.0908 | 73.3662 | 0.088247 | 4.8654 | 1.0064 | 0.20058 | 2.3851 | 0.074664
| K5 | 36 | 44.3167 | 0 | 0.00085608 | 0.0098809 | 0.00055005 | 0.0011669 | 0.017616 | 0.0057891 | 0.0092154 | 0.0052627 | 0.0091319 | 0.05329 | 0.018086 | 0.057623 | 0.0047575 | 0.019911 | 0.0076926 | 0.011264 | 3.8703 | 0.13902 | 14.189 | 73.0392 | 0.059845 | 4.8316 | 0.99866 | 0.19445 | 2.371 | 0.074888
| K5 | 37 | 44.4 | 0 | 0.00167 | 0.0098853 | 0.00053305 | 0.001357 | 0.017282 | 0.005527 | 0.0096146 | 0.0056297 | 0.0084579 | 0.052994 | 0.017761 | 0.047829 | 0.0019241 | 0.022363 | 0.0078076 | 0.011402 | 3.9 | 0.11048 | 14.0602 | 73.099 | 0.11444 | 4.8531 | 0.99037 | 0.20279 | 2.3734 | 0.074062
| K5 | 38 | 44.9 | 0 | 0.0015102 | 0.010154 | 0.00064878 | 0.0015864 | 0.017317 | 0.0057272 | 0.0099065 | 0.0058631 | 0.0081685 | 0.052765 | 0.01822 | 0.048315 | 0.00089755 | 0.022343 | 0.0074788 | 0.010876 | 3.9806 | 0.2384 | 14.0678 | 72.9991 | 0.059036 | 4.8126 | 0.98733 | 0.20117 | 2.359 | 0.073228
| K5 | 39 | 45.4 | 0 | 0.0014654 | 0.010018 | 0.00054497 | 0.0019362 | 0.017845 | 0.0064718 | 0.0093387 | 0.0059645 | 0.0087529 | 0.053474 | 0.017606 | 0.050614 | 0.0035163 | 0.024068 | 0.0069246 | 0.010624 | 3.5243 | 0.11334 | 14.0721 | 73.4479 | 0.087042 | 4.8632 | 0.99564 | 0.20088 | 2.3915 | 0.074946
| K5 | 40 | 45.9 | 0 | 0.0012749 | 0.0099109 | 0.00069492 | 0.0012859 | 0.017312 | 0.0054908 | 0.0098007 | 0.0055975 | 0.0084713 | 0.053114 | 0.018199 | 0.051696 | 0.004292 | 0.021642 | 0.0082993 | 0.011612 | 3.9498 | 0.2388 | 14.1317 | 72.9594 | 0.062361 | 4.8043 | 0.98247 | 0.20112 | 2.3675 | 0.073846
| K5 | 41 | 46.4 | 0 | 0.0016217 | 0.0095811 | 0.0002835 | 0.0016025 | 0.01722 | 0.0061805 | 0.0091062 | 0.0060819 | 0.0092302 | 0.053172 | 0.018051 | 0.04422 | 0.00099216 | 0.025455 | 0.0067599 | 0.010443 | 3.8917 | 0.18245 | 14.1216 | 73.0575 | 0.080576 | 4.8154 | 0.99053 | 0.20047 | 2.3674 | 0.072354
| K5 | 42 | 46.9 | 0 | 0.00074449 | 0.0095926 | 0.00048552 | 0.0017884 | 0.017418 | 0.0052235 | 0.0095136 | 0.0060352 | 0.0092566 | 0.053549 | 0.01757 | 0.053599 | 0.002857 | 0.023505 | 0.0077746 | 0.011461 | 3.7133 | 0.16293 | 14.0191 | 73.2619 | 0.079368 | 4.8673 | 1.0051 | 0.19992 | 2.385 | 0.075753
| K5 | 43 | 47.4 | 0 | 0.00059427 | 0.010853 | 0.00063173 | 0.0014897 | 0.017174 | 0.0058625 | 0.0093274 | 0.0057038 | 0.008351 | 0.053726 | 0.018936 | 0.053659 | 0.0027771 | 0.018485 | 0.0070386 | 0.010306 | 3.498 | 0.16096 | 14.0461 | 73.5035 | 0.064036 | 4.8929 | 0.95895 | 0.20095 | 2.3767 | 0.073002
| K5 | 44 | 47.9 | 0 | 8.4074e-05 | 0.010608 | 0.00089194 | 0.0015023 | 0.017655 | 0.0054987 | 0.0091713 | 0.0059812 | 0.0083434 | 0.052399 | 0.017374 | 0.055099 | 0.006354 | 0.021534 | 0.0088264 | 0.011235 | 3.8136 | 0.17972 | 14.1379 | 73.1098 | 0.072537 | 4.8723 | 0.95916 | 0.19416 | 2.3547 | 0.073564
| K5 | 45 | 48.4 | 0 | 0 | 0.010312 | 0.0004722 | 0.0017159 | 0.017125 | 0.0059324 | 0.0089313 | 0.0057416 | 0.0085621 | 0.052172 | 0.018385 | 0.060267 | 0.0039721 | 0.020982 | 0.0070637 | 0.0099643 | 3.8601 | 0.20808 | 13.9662 | 73.2143 | 0.06711 | 4.8724 | 0.96328 | 0.19286 | 2.3499 | 0.074183

## Table S2 Published Standards
Published data for standards and knowns, all data recalculated for MnO and Fe2O3(total) values and renormalised to 100% for the elements of interest.
Trace element data quoted in weight percent to facilitate rapid and easy comparison with M4 Tornado data output files.
|  | BHVO-2 |  |  | AGV-2 |  |  | GSP-2 |  |  | GOR-132 |  |  | T1-G |  |  | K5 |  |  | 00L-31A
| Na2O | 2.2108 | ± | 0.079811 | 4.2389 | ± | 0.13186 | 2.8094 | ± | 0.091267 | 0.82697 | ± | 0.040028 | 3.1301 | ± | 0.090138 | 4.9305 | ± | 0.049995 | 3.738 | ± | 0.037903
| MgO | 7.1999 | ± | 0.11972 | 1.8109 | ± | 0.030429 | 0.97017 | ± | 0.030422 | 22.3181 | ± | 0.20014 | 3.7501 | ± | 0.060092 | 0.16967 | ± | 0.020998 | 0 | ± | 0
| Al2O3 | 13.4438 | ± | 0.19953 | 17.1073 | ± | 0.213 | 15.0578 | ± | 0.20282 | 10.9598 | ± | 0.10007 | 17.1003 | ± | 0.20031 | 13.2745 | ± | 0.009999 | 12.749 | ± | 0.0096032
| SiO2 | 49.6924 | ± | 0.59858 | 59.9917 | ± | 0.71002 | 67.3054 | ± | 0.81127 | 45.3337 | ± | 0.50034 | 58.6011 | ± | 0.40061 | 73.6787 | ± | 0.069993 | 76.0655 | ± | 0.07226
| P2O5 | 0.26888 | ± | 0.019953 | 0.4856 | ± | 0.020286 | 0.29307 | ± | 0.020282 | 0.035868 | ± | 0.012008 | 0.168 | ± | 0.0060092 | 0.019962 | ± | 0 | 0.48843 | ± | 0
| K2O | 0.51784 | ± | 0.0099764 | 2.9136 | ± | 0.11157 | 5.437 | ± | 0.14197 | 0.030687 | ± | 0.0012008 | 1.96 | ± | 0.040061 | 3.9125 | ± | 0.019998 | 5.4525 | ± | 0.027869
| CaO | 11.3526 | ± | 0.19953 | 5.2607 | ± | 0.13186 | 2.1222 | ± | 0.060845 | 8.4191 | ± | 0.12008 | 7.1001 | ± | 0.050077 | 1.0081 | ± | 0.019998 | 0.40869 | ± | 0.0081075
| TiO2 | 2.7186 | ± | 0.039906 | 1.0622 | ± | 0.22315 | 0.66699 | ± | 0.020282 | 0.30488 | ± | 0.016011 | 0.75501 | ± | 0.032049 | 0.2096 | ± | 0 | 0.10965 | ± | 0
| MnO | 0.16586 | ± | 0.0051523 | 0.10058 | ± | 0.0026192 | 0.041754 | ± | 0.0026186 | 0.15344 | ± | 0.0070048 | 0.127 | ± | 0.0040061 | 0.069866 | ± | 0 | 0.10965 | ± | 0
| Fe2O3t | 12.2488 | ± | 0.19953 | 6.768 | ± | 0.10143 | 4.9519 | ± | 0.16225 | 11.1833 | ± | 0.11121 | 7.157 | ± | 0.055651 | 2.5451 | ± | 0.009999 | 0.81737 | ± | 0.0032112
| (Total) | 99.8194 |  |  | 99.7394 |  |  | 99.6558 |  |  | 99.5659 |  |  | 99.8487 |  |  | 99.8185 |  |  | 99.9388
| S | nd |  |  | nd |  |  | nd |  |  | 0.00017934 | ± | nd | 0.00026 | ± | nd | 0.0025751 | ± | 0.00076992 | nd
| V | 0.031568 | ± | 0.0010974 | 0.01214 | ± | 0.00050715 | 0.0052551 | ± | 0.00040563 | 0.021322 | ± | 0.0018012 | 0.019 | ± | 0.0011017 | 0.00017966 | ± | 3.4997e-05 | 0.0011962 | ± | 0.00023301
| Cr | 0.027883 | ± | 0.0018955 | 0.0017198 | ± | 0.00020286 | 0.0020212 | ± | 0.00060845 | 0.25188 | ± | 0.010807 | 0.00209 | ± | 0.00010015 | 0.001068 | ± | 0.00011299 | 0.00039872 | ± | 4.2184e-05
| Co | 0.0044813 | ± | 0.00029929 | 0.0016187 | ± | 0.00010143 | 0.00073773 | ± | 8.1127e-05 | 0.0092361 | ± | 0.00042029 | 0.00189 | ± | 0.0001302 | 0.00052899 | ± | 0.00016998 | nd
| Ni | 0.01185 | ± | 0.00069835 | 0.0019222 | ± | 0.00030429 | 0.001718 | ± | 0.00020282 | 0.11827 | ± | 0.0021014 | 0.00106 | ± | 0.00017026 | 0.00031939 | ± | 0.00011999 | nd
| Cu | 0.012647 | ± | 0.00069835 | 0.0053618 | ± | 0.00040572 | 0.0043455 | ± | 0.00040563 | 0.020425 | ± | 0.00090062 | 0.00188 | ± | 0.00022034 | 0.00092822 | ± | 7.7992e-05 | nd
| Zn | 0.010257 | ± | 0.00059858 | 0.0087003 | ± | 0.00081145 | 0.012127 | ± | 0.0010141 | 0.0076519 | ± | 0.00056039 | 0.0074001 | ± | 0.00090138 | 0.013275 | ± | 0.00055994 | 0.0057814 | ± | 0.00024387
| Ga | 0.002161 | ± | 8.9788e-05 | 0.0020233 | ± | 0.00010143 | 0.0022233 | ± | 0.00020282 | 0.0010362 | ± | 4.0028e-05 | 0.00194 | ± | 0.00011017 | 0.0024753 | ± | 1.3999e-05 | 0.0019936 | ± | 1.1275e-05
| Rb | 0.00097592 | ± | 9.9764e-05 | 0.00694 | ± | 0.00023329 | 0.024759 | ± | 0.00070986 | 0.00020923 | ± | 2.0014e-06 | 0.0079701 | ± | 0.00054083 | 0.0095417 | ± | 7.0993e-05 | 0.020534 | ± | 0.00015278
| Sr | 0.038738 | ± | 0.0022946 | 0.066568 | ± | 0.0017243 | 0.024254 | ± | 0.0010141 | 0.0015244 | ± | 0.0001501 | 0.028401 | ± | 0.00080123 | 0.005759 | ± | 0.00015598 | 0 | ± | 0
| Y | 0.0025892 | ± | 0.00019953 | 0.0020233 | ± | 0.00010143 | 0.0028297 | ± | 0.00020282 | 0.0012853 | ± | 0.00013009 | 0.00239 | ± | 0.00011017 | 0.010739 | ± | 0.00017698 | 0.0022926 | ± | 3.7782e-05
| Zr | 0.017128 | ± | 0.0010974 | 0.023268 | ± | 0.00040572 | 0.055583 | ± | 0.0030422 | 0.00098638 | ± | 5.0034e-05 | 0.0144 | ± | 0.00060092 | 0.043157 | ± | 0.00069993 | 0.013058 | ± | 0.00021178
| Nb | 0.0017925 | ± | 0.00019953 | 0.0015175 | ± | 0.00010143 | 0.0027286 | ± | 0.00020282 | 7.2733e-06 | ± | 1.0007e-06 | 0.00088702 | ± | 0.00016525 | 0.012536 | ± | 2.7997e-05 | 0.0030901 | ± | 6.9012e-06
| Ba | 0.012946 | ± | 0.0012969 | 0.11533 | ± | 0.0032458 | 0.13542 | ± | 0.004462 | 8.1202e-05 | ± | 6.1042e-06 | 0.038801 | ± | 0.0019029 | 0.076992 | ± | 0.0025997 | 0.00079744 | ± | 2.6926e-05
| La | 0.0014938 | ± | 9.9764e-05 | 0.0038443 | ± | 0.00010143 | 0.018191 | ± | 0.0012169 | 8.3892e-06 | ± | 6.5045e-07 | 0.0070401 | ± | 0.00022034 | nd |  |  | 0.0036881 | ± | nd
| Ce | 0.0037842 | ± | 0.00019953 | 0.0068793 | ± | 0.00030429 | 0.041434 | ± | 0.0030422 | 3.9156e-05 | ± | 1.8012e-06 | 0.0127 | ± | 0.00080123 | nd |  |  | 0.0066785 | ± | nd
| Ta | 0.00013942 | ± | nd | 9.0038e-05 | ± | 8.1145e-06 | 0 | ± | 0 | 3.0887e-06 | ± | 4.0028e-07 | 4.6401e-05 | ± | 2.1032e-06 | nd |  |  | 0 | ± | nd
| Th | 0.0001195 | ± | 2.9929e-05 | 0.00061712 | ± | 6.0858e-05 | 0.010611 | ± | 0.00081127 | 8.9671e-07 | ± | 6.0041e-07 | 0.0031201 | ± | 0.00015023 | 0.0013873 | ± | 5.6994e-05 | 0.0016946 | ± | 6.9615e-05
| Total | 100 |  |  | 100 |  |  | 100 |  |  | 100 |  |  | 100 |  |  | 100 |  |  | 100

## Table S3 Analysis time
Variation in gross intensity and error due to counting statistics for each element measured in standard GSP-2
|  |  | Gross Intensity (gross counts, ROI): |  |  |  |  | Standard deviation (square root of intensity): |  |  |  |  | Relative % error
|  | Analysis time | 600s | 300s | 120s | 60s | 30s | 600s | 300s | 120s | 60s | 30s | 600s | 300s | 120s | 60s | 30s
| z | Element
| 11 | Na | 23169 | 11816 | 4765 | 2342 | 1199 | 152.2137 | 108.7014 | 69.029 | 48.3942 | 34.6266 | 0.65697 | 0.91995 | 1.4487 | 2.0664 | 2.888
| 12 | Mg | 28094 | 14174 | 5773 | 2793 | 1385 | 167.6126 | 119.0546 | 75.9803 | 52.8488 | 37.2156 | 0.59661 | 0.83995 | 1.3161 | 1.8922 | 2.687
| 13 | Al | 240800 | 120701 | 48166 | 23975 | 12046 | 490.7138 | 347.4205 | 219.4675 | 154.8386 | 109.7543 | 0.20378 | 0.28784 | 0.45565 | 0.64583 | 0.91113
| 14 | Si | 1680018 | 839235 | 334116 | 166925 | 83803 | 1296.1551 | 916.0977 | 578.0277 | 408.5646 | 289.4875 | 0.077151 | 0.10916 | 0.173 | 0.24476 | 0.34544
| 15 | P | 27790 | 13924 | 5541 | 2757 | 1446 | 166.7033 | 118 | 74.4379 | 52.5071 | 38.0263 | 0.59987 | 0.84746 | 1.3434 | 1.9045 | 2.6298
| 19 | K | 864082 | 430399 | 171523 | 86213 | 42963 | 929.5601 | 656.048 | 414.1534 | 293.6205 | 207.2752 | 0.10758 | 0.15243 | 0.24146 | 0.34058 | 0.48245
| 20 | Ca | 476535 | 237819 | 94602 | 47796 | 23983 | 690.3151 | 487.6669 | 307.5744 | 218.623 | 154.8645 | 0.14486 | 0.20506 | 0.32512 | 0.45741 | 0.64573
| 22 | Ti | 254582 | 127108 | 50664 | 25442 | 12581 | 504.5612 | 356.5221 | 225.0866 | 159.5055 | 112.1651 | 0.19819 | 0.28049 | 0.44427 | 0.62694 | 0.89154
| 23 | V | 91659 | 45718 | 18104 | 8987 | 4473 | 302.7524 | 213.8177 | 134.5511 | 94.7998 | 66.8805 | 0.3303 | 0.46769 | 0.74321 | 1.0549 | 1.4952
| 24 | Cr | 52166 | 26083 | 10323 | 5202 | 2642 | 228.3988 | 161.5023 | 101.6022 | 72.1249 | 51.4004 | 0.43783 | 0.61919 | 0.98423 | 1.3865 | 1.9455
| 25 | Mn | 98684 | 49123 | 19772 | 9853 | 4823 | 314.1401 | 221.6371 | 140.6129 | 99.2623 | 69.4478 | 0.31833 | 0.45119 | 0.71117 | 1.0074 | 1.4399
| 26 | Fe | 4953178 | 2471390 | 987874 | 493601 | 247289 | 2225.5736 | 1572.0655 | 993.9185 | 702.5674 | 497.2816 | 0.044932 | 0.063611 | 0.10061 | 0.14234 | 0.20109
| 27 | Co | 469750 | 234200 | 93338 | 46700 | 23147 | 685.3831 | 483.9421 | 305.5127 | 216.1018 | 152.1414 | 0.1459 | 0.20664 | 0.32732 | 0.46274 | 0.65728
| 28 | Ni | 60819 | 30029 | 12451 | 6203 | 3034 | 246.6151 | 173.2888 | 111.584 | 78.7591 | 55.0818 | 0.40549 | 0.57707 | 0.89619 | 1.2697 | 1.8155
| 29 | Cu | 79861 | 40173 | 16013 | 7972 | 3979 | 282.5969 | 200.432 | 126.5425 | 89.2861 | 63.0793 | 0.35386 | 0.49892 | 0.79025 | 1.12 | 1.5853
| 30 | Zn | 111080 | 55679 | 22108 | 11096 | 5544 | 333.2867 | 235.964 | 148.6876 | 105.3376 | 74.458 | 0.30004 | 0.42379 | 0.67255 | 0.94933 | 1.343
| 31 | Ga | 102296 | 51062 | 20441 | 10386 | 5092 | 319.8375 | 225.969 | 142.972 | 101.9117 | 71.3583 | 0.31266 | 0.44254 | 0.69944 | 0.98124 | 1.4014
| 37 | Rb | 113036 | 56400 | 23001 | 11407 | 5834 | 336.2083 | 237.4868 | 151.6608 | 106.8036 | 76.3806 | 0.29743 | 0.42108 | 0.65937 | 0.9363 | 1.3092
| 38 | Sr | 100819 | 50758 | 20227 | 10216 | 5151 | 317.5201 | 225.2954 | 142.2217 | 101.0742 | 71.7705 | 0.31494 | 0.44386 | 0.70313 | 0.98937 | 1.3933
| 39 | Y | 64881 | 32768 | 13041 | 6524 | 3242 | 254.7175 | 181.0193 | 114.1972 | 80.7713 | 56.9386 | 0.39259 | 0.55243 | 0.87568 | 1.2381 | 1.7563
| 40 | Zr | 115144 | 58179 | 23434 | 11682 | 5882 | 339.3287 | 241.2032 | 153.0817 | 108.0833 | 76.6942 | 0.2947 | 0.41459 | 0.65325 | 0.92521 | 1.3039
| 41 | Nb | 42104 | 21191 | 8289 | 4259 | 2054 | 205.1926 | 145.5713 | 91.0439 | 65.261 | 45.3211 | 0.48735 | 0.68695 | 1.0984 | 1.5323 | 2.2065
| 56 | Ba | 236931 | 117989 | 47316 | 23672 | 11659 | 486.7556 | 343.4953 | 217.5224 | 153.8571 | 107.9768 | 0.20544 | 0.29112 | 0.45972 | 0.64995 | 0.92612
| 57 | La | 100143 | 50042 | 19728 | 10003 | 4970 | 316.4538 | 223.7007 | 140.4564 | 100.015 | 70.4982 | 0.316 | 0.44703 | 0.71196 | 0.99985 | 1.4185
| 58 | Ce | 84772 | 42335 | 16837 | 8402 | 4122 | 291.1563 | 205.7547 | 129.7575 | 91.6624 | 64.2028 | 0.34346 | 0.48602 | 0.77067 | 1.091 | 1.5576
| 90 | Th | 87193 | 43576 | 17492 | 8693 | 4402 | 295.2846 | 208.7487 | 132.2573 | 93.2363 | 66.3476 | 0.33866 | 0.47905 | 0.7561 | 1.0725 | 1.5072

## Table S4 GSP-2
Non-calibrated concentrations (wt%): GSP-2 | [19 empty] | Concentration (wt%) after manual recalibration of spectra using the 0 and Cu peaks
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) |  | Calibration factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | % | [6 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |  | Mean (n=10) | ± 1 σ
| Na2O | 3.0639 | 2.9751 | 3.2902 | 2.9402 | 3.1308 | 2.9453 | 3.1388 | 2.9002 | 2.8841 | 3.1883 | 3.0457 | 0.13738 | 4.5106 | 0.23625 | 8.409 |  | 0.92243 |  | Na2O | 3.2506 | 3.1151 | 3.4892 | 3.1997 | 3.3292 | 3.155 | 3.3597 | 3.0709 | 3.1021 | 3.3966 |  | 3.2468 | 0.14166
| MgO | 0.49738 | 0.49178 | 0.48386 | 0.47837 | 0.4611 | 0.52927 | 0.48019 | 0.52246 | 0.48949 | 0.45841 | 0.48923 | 0.022973 | 4.6957 | -0.48094 | -49.5726 |  | 1.983 |  | MgO | 0.7767 | 0.64853 | 0.73113 | 0.73851 | 0.68372 | 0.79694 | 0.72182 | 0.7793 | 0.75289 | 0.71298 |  | 0.73425 | 0.045514
| Al2O3 | 15.2494 | 15.308 | 15.3555 | 15.3127 | 15.3342 | 15.3278 | 15.3599 | 15.3126 | 15.2355 | 15.2524 | 15.3048 | 0.044408 | 0.29016 | 0.24699 | 1.6403 |  | 0.98386 |  | Al2O3 | 15.2087 | 14.9329 | 15.2021 | 15.2437 | 15.2128 | 15.2568 | 15.2197 | 15.2075 | 15.1187 | 15.1818 |  | 15.1785 | 0.09397
| SiO2 | 65.7564 | 65.8563 | 65.5832 | 65.85 | 65.7315 | 65.771 | 65.6433 | 65.7931 | 65.888 | 65.658 | 65.7531 | 0.10034 | 0.1526 | -1.5523 | -2.3064 |  | 1.0236 |  | SiO2 | 67.2632 | 66.9097 | 66.6624 | 67.2006 | 66.9168 | 67.0287 | 66.8266 | 67.1805 | 67.264 | 67.101 |  | 67.0354 | 0.2026
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.29307 | -100 |  | nd |  | P2O5 | 0.015351 | 0 | 0 | 0.017237 | 0 | 0.010942 | 0 | 0 | 0.0056732 | 0.0084462 |  | 0.0057648 | 0.0068609
| K2O | 7.4684 | 7.4376 | 7.4059 | 7.4342 | 7.4124 | 7.4467 | 7.4324 | 7.4844 | 7.5059 | 7.4849 | 7.4513 | 0.03323 | 0.44596 | 2.0143 | 37.048 |  | 0.72967 |  | K2O | 6.749 | 7.0421 | 6.8365 | 6.6962 | 6.8031 | 6.8376 | 6.8147 | 6.7841 | 6.8269 | 6.7874 |  | 6.8178 | 0.090027
| CaO | 2.0187 | 2.0163 | 2.0003 | 2.0255 | 2.0105 | 2.0174 | 2.003 | 2.0207 | 2.0229 | 2.0107 | 2.0146 | 0.008308 | 0.41239 | -0.10764 | -5.072 |  | 1.0534 |  | CaO | 1.8067 | 1.8995 | 1.8206 | 1.8161 | 1.8235 | 1.8072 | 1.8167 | 1.8258 | 1.8154 | 1.7946 |  | 1.8226 | 0.028547
| TiO2 | 0.6339 | 0.63649 | 0.62606 | 0.63176 | 0.63475 | 0.63259 | 0.62371 | 0.62472 | 0.62289 | 0.61908 | 0.62859 | 0.0059891 | 0.95278 | -0.038397 | -5.7568 |  | 1.0611 |  | TiO2 | 0.44766 | 0.47358 | 0.46164 | 0.45566 | 0.46578 | 0.42767 | 0.45372 | 0.44806 | 0.4318 | 0.42041 |  | 0.4486 | 0.017262
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0052551 | -100 |  | nd |  | V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
| Cr | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0020212 | -100 |  | nd |  | Cr | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
| MnO | 0.033788 | 0.033507 | 0.034065 | 0.034533 | 0.034676 | 0.033563 | 0.034016 | 0.033723 | 0.034161 | 0.034376 | 0.034041 | 0.00040157 | 1.1797 | -0.0077129 | -18.4725 |  | 1.2266 |  | MnO | 0.033518 | 0.035562 | 0.035503 | 0.035984 | 0.036123 | 0.03385 | 0.035367 | 0.035504 | 0.034712 | 0.034351 |  | 0.035047 | 0.0008936
| Fe2O3t | 5.1442 | 5.1444 | 5.1063 | 5.1432 | 5.1254 | 5.151 | 5.1234 | 5.1565 | 5.1611 | 5.1394 | 5.1395 | 0.016691 | 0.32476 | 0.1876 | 3.7885 |  | 0.9635 |  | Fe2O3 | 4.1019 | 4.6255 | 4.4162 | 4.2539 | 4.3815 | 4.2666 | 4.3607 | 4.3038 | 4.2624 | 4.183 |  | 4.3156 | 0.14336
| Co | 0.01901 | 0.018769 | 0.018421 | 0.019119 | 0.018372 | 0.018852 | 0.018876 | 0.019735 | 0.020355 | 0.019397 | 0.019091 | 0.00060434 | 3.1656 | 0.018353 | 2487.755 |  | 0.038644 |  | Co | 0.053957 | 0.057206 | 0.056214 | 0.051851 | 0.051674 | 0.059323 | 0.053296 | 0.054928 | 0.059989 | 0.061074 |  | 0.055951 | 0.0033787
| Ni | 0.00093887 | 0.0012517 | 0.0011564 | 0.0012702 | 0.00090145 | 0.00088083 | 0.0010798 | 0.0008645 | 0.0010069 | 0.00075792 | 0.0010109 | 0.00017342 | 17.1558 | -0.00070715 | -41.1614 |  | 1.6996 |  | Ni | 0.0011249 | 0.0012685 | 0.0013668 | 0.0013722 | 0.0011881 | 0.001203 | 0.0012876 | 0.0011412 | 0.0011869 | 0.0010854 |  | 0.0012225 | 9.8534e-05
| Cu | 0.0051419 | 0.0051855 | 0.0050372 | 0.0053053 | 0.0051087 | 0.0052972 | 0.0053145 | 0.0054229 | 0.0051302 | 0.0052178 | 0.0052161 | 0.00011754 | 2.2534 | 0.00087056 | 20.0334 |  | 0.8331 |  | Cu | 0.0035086 | 0.0040439 | 0.0039874 | 0.003631 | 0.0039418 | 0.003871 | 0.0040238 | 0.0038737 | 0.0035581 | 0.0037923 |  | 0.0038232 | 0.00019492
| Zn | 0.013746 | 0.013823 | 0.013348 | 0.013908 | 0.013457 | 0.013879 | 0.013981 | 0.013912 | 0.014155 | 0.013353 | 0.013756 | 0.00027803 | 2.0211 | 0.0016292 | 13.4345 |  | 0.88157 |  | Zn | 0.0095141 | 0.011189 | 0.010477 | 0.0097161 | 0.010272 | 0.010004 | 0.010666 | 0.0099887 | 0.01001 | 0.0093315 |  | 0.010117 | 0.00055627
| Ga | 0.0030809 | 0.0031057 | 0.0029191 | 0.0028403 | 0.0033098 | 0.003108 | 0.0029692 | 0.0031554 | 0.0029607 | 0.0029942 | 0.0030443 | 0.00013514 | 4.4392 | 0.00082103 | 36.9284 |  | 0.73031 |  | Ga | 0.0018404 | 0.0019555 | 0.0019716 | 0.0016388 | 0.0023181 | 0.0019946 | 0.0019462 | 0.0019217 | 0.0018771 | 0.0017892 |  | 0.0019253 | 0.00017409
| Rb | 0.027489 | 0.02801 | 0.027775 | 0.02844 | 0.028037 | 0.028398 | 0.027802 | 0.028265 | 0.02841 | 0.028433 | 0.028106 | 0.00033639 | 1.1969 | 0.0033464 | 13.5157 |  | 0.88094 |  | Rb | 0.020222 | 0.023133 | 0.022279 | 0.021537 | 0.022085 | 0.021396 | 0.021869 | 0.021541 | 0.02126 | 0.020783 |  | 0.021611 | 0.00080599
| Sr | 0.028037 | 0.027768 | 0.028022 | 0.028048 | 0.028029 | 0.027983 | 0.027894 | 0.028388 | 0.028846 | 0.028622 | 0.028164 | 0.0003424 | 1.2157 | 0.0039095 | 16.1188 |  | 0.86119 |  | Sr | 0.020508 | 0.022607 | 0.021584 | 0.021085 | 0.021452 | 0.020354 | 0.021338 | 0.021132 | 0.020778 | 0.020158 |  | 0.0211 | 0.00071365
| Y | 0.0022287 | 0.0022438 | 0.0021319 | 0.0023047 | 0.0018319 | 0.0021592 | 0.001887 | 0.0022628 | 0.002127 | 0.0019816 | 0.0021159 | 0.00016319 | 7.7126 | -0.00071379 | -25.2252 |  | 1.3373 |  | Y | 0.0019463 | 0.0021805 | 0.0020465 | 0.0018637 | 0.0018434 | 0.001807 | 0.0015313 | 0.0017841 | 0.0019428 | 0.001651 |  | 0.0018597 | 0.0001863
| Zr | 0.051062 | 0.051147 | 0.050736 | 0.051836 | 0.051357 | 0.052803 | 0.052913 | 0.052946 | 0.053901 | 0.054062 | 0.052276 | 0.0012076 | 2.3101 | -0.0033063 | -5.9484 |  | 1.0632 |  | Zr | 0.038658 | 0.043953 | 0.040866 | 0.039632 | 0.041314 | 0.040156 | 0.042052 | 0.040998 | 0.041101 | 0.040437 |  | 0.040917 | 0.0014247
| Nb | 0.0044531 | 0.0037648 | 0.0042445 | 0.0045086 | 0.0043898 | 0.0043107 | 0.0040099 | 0.00402 | 0.0044318 | 0.0039739 | 0.0042107 | 0.00025193 | 5.983 | 0.0014821 | 54.3179 |  | 0.64801 |  | Nb | 0.0025789 | 0.0027478 | 0.002783 | 0.0026369 | 0.0029913 | 0.0027811 | 0.002521 | 0.0022828 | 0.0028551 | 0.0025383 |  | 0.0026716 | 0.00020208
| Ba | 0.16541 | 0.14958 | 0.15569 | 0.1731 | 0.16345 | 0.18106 | 0.20025 | 0.18746 | 0.19234 | 0.18866 | 0.1757 | 0.016868 | 9.6001 | 0.040282 | 29.7459 |  | 0.77074 |  | Ba | 0.12698 | 0.14396 | 0.13816 | 0.12301 | 0.1396 | 0.15042 | 0.16544 | 0.13798 | 0.15218 | 0.14454 |  | 0.14223 | 0.012278
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.018191 | -100 |  | nd |  | La | 0.056055 | 0.055654 | 0.056407 | 0.054893 | 0.058466 | 0.054277 | 0.059833 | 0.059338 | 0.057808 | 0.062225 |  | 0.057496 | 0.0024956
| Ce | 0.01644 | 0.020231 | 0.021388 | 0.019704 | 0.019861 | 0.018194 | 0.0076241 | 0.017386 | 0.010867 | 0.011357 | 0.016305 | 0.0047095 | 28.8838 | -0.025129 | -60.6482 |  | 2.5412 |  | Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
| Th | 0.014822 | 0.015408 | 0.014127 | 0.014539 | 0.015602 | 0.015287 | 0.014255 | 0.013984 | 0.014862 | 0.014278 | 0.014716 | 0.00057167 | 3.8846 | 0.0041052 | 38.6875 |  | 0.72105 |  | Th | 0.0097496 | 0.012685 | 0.010166 | 0.0095563 | 0.011819 | 0.011154 | 0.010893 | 0.0097932 | 0.01082 | 0.009923 |  | 0.010656 | 0.0010195
Type-calibrated concentrations (wt%): GSP-2
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 2.8993 | 2.8144 | 3.1144 | 2.7841 | 2.9657 | 2.784 | 2.9727 | 2.7424 | 2.7303 | 3.0222 | 2.8895 | 0.14645 | 5.0684 | 0.080025 | 2.8484
| MgO | 1.4686 | 1.4518 | 1.429 | 1.4134 | 1.3629 | 1.561 | 1.4189 | 1.5415 | 1.4459 | 1.3557 | 1.441 | 0.074907 | 5.1981 | 0.47088 | 48.5359
| Al2O3 | 15.96 | 16.0154 | 16.069 | 16.0293 | 16.0515 | 16.0346 | 16.0804 | 16.0221 | 15.949 | 15.9742 | 16.0263 | 0.045001 | 0.28079 | 0.96846 | 6.4316
| SiO2 | 66.6447 | 66.7211 | 66.4656 | 66.7627 | 66.6452 | 66.6236 | 66.5656 | 66.6648 | 66.8019 | 66.6111 | 66.6426 | 0.10619 | 0.15934 | -0.66282 | -0.98479
| P2O5 | 0.033597 | 0.03362 | 0.033417 | 0.033556 | 0.033688 | 0.033557 | 0.033602 | 0.033555 | 0.033853 | 0.033643 | 0.033609 | 0.00012656 | 0.37655 | -0.25946 | -88.5322
| K2O | 5.6392 | 5.6138 | 5.5926 | 5.617 | 5.6017 | 5.6177 | 5.6167 | 5.6481 | 5.6701 | 5.6605 | 5.6281 | 0.028115 | 0.49955 | 0.19108 | 3.5144
| CaO | 2.068 | 2.0651 | 2.0501 | 2.0767 | 2.0622 | 2.0648 | 2.0542 | 2.0689 | 2.0733 | 2.0635 | 2.0642 | 0.0089905 | 0.43554 | -0.058039 | -2.7348
| TiO2 | 0.59346 | 0.59679 | 0.5874 | 0.59109 | 0.59645 | 0.59202 | 0.58562 | 0.58726 | 0.58542 | 0.58147 | 0.58834 | 0.0046618 | 0.79237 | -0.07865 | -11.7918
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0052551 | -100
| Cr | 0.00018667 | 0.00047673 | 0.00033127 | 0.00079807 | 0.00055843 | 0.00063117 | 0.0007623 | 0.00077362 | 0.0005272 | 0.00041532 | 0.00059967 | 0.00017322 | 28.8859 | -0.0014215 | -70.3307
| MnO | 0.027403 | 0.027613 | 0.027724 | 0.028121 | 0.028313 | 0.027162 | 0.027552 | 0.027455 | 0.027418 | 0.02776 | 0.027688 | 0.00037937 | 1.3702 | -0.014065 | -33.6868
| Fe2O3 | 4.4751 | 4.475 | 4.4444 | 4.4753 | 4.4639 | 4.4722 | 4.4565 | 4.4768 | 4.4881 | 4.4762 | 4.4692 | 0.013714 | 0.30687 | -0.48272 | -9.7481
| Co | 0.0010907 | 0.0010773 | 0.0010573 | 0.0010971 | 0.0010551 | 0.0010789 | 0.0010828 | 0.0011301 | 0.0011677 | 0.0011141 | 0.0010979 | 3.8242e-05 | 3.4833 | 0.00036014 | 48.8176
| Ni | 0.0015567 | 0.002076 | 0.0019193 | 0.0021077 | 0.0014971 | 0.0014577 | 0.0017909 | 0.0014302 | 0.0016689 | 0.0012582 | 0.0016413 | 0.0002838 | 17.2915 | -7.6751e-05 | -4.4675
| Cu | 0.0034774 | 0.0035069 | 0.0034087 | 0.0035888 | 0.0034596 | 0.0035747 | 0.0035941 | 0.0036597 | 0.003468 | 0.0035333 | 0.0035359 | 8.4194e-05 | 2.3811 | -0.00080968 | -18.6324
| Zn | 0.011248 | 0.011312 | 0.010929 | 0.011384 | 0.011027 | 0.011332 | 0.01144 | 0.011359 | 0.011577 | 0.01094 | 0.011249 | 0.00024739 | 2.1993 | -0.00087848 | -7.244
| Ga | 0.0017279 | 0.0017419 | 0.0016382 | 0.0015934 | 0.0018588 | 0.0017393 | 0.0016652 | 0.0017658 | 0.0016597 | 0.0016815 | 0.0017002 | 8.4057e-05 | 4.9438 | -0.00052306 | -23.5261
| Rb | 0.017927 | 0.018269 | 0.018126 | 0.018551 | 0.01831 | 0.018476 | 0.018127 | 0.018388 | 0.018516 | 0.018566 | 0.018382 | 0.00017929 | 0.97534 | -0.006377 | -25.7559
| Sr | 0.021688 | 0.021481 | 0.021691 | 0.021699 | 0.021711 | 0.021593 | 0.021571 | 0.021904 | 0.022298 | 0.022165 | 0.021829 | 0.00027013 | 1.2375 | -0.0024251 | -9.9988
| Y | 0.0029681 | 0.0029881 | 0.0028406 | 0.0030696 | 0.00244 | 0.002867 | 0.0025093 | 0.003005 | 0.002829 | 0.0026394 | 0.002775 | 0.00022568 | 8.1326 | -5.4674e-05 | -1.9322
| Zr | 0.031214 | 0.031269 | 0.031035 | 0.031689 | 0.031437 | 0.032197 | 0.032334 | 0.032282 | 0.032924 | 0.033085 | 0.032123 | 0.00070555 | 2.1964 | -0.02346 | -42.2069
| Nb | 0.0016765 | 0.0014179 | 0.0015992 | 0.0016975 | 0.001655 | 0.001619 | 0.0015095 | 0.0015099 | 0.0016673 | 0.0014982 | 0.0015944 | 7.9147e-05 | 4.964 | -0.0011342 | -41.5658
| Ba | 0.081754 | 0.075177 | 0.078447 | 0.077669 | 0.078606 | 0.078439 | 0.078245 | 0.076753 | 0.082317 | 0.083284 | 0.07922 | 0.002303 | 2.9071 | -0.056199 | -41.5003
| Th | 0.014 | 0.014551 | 0.013353 | 0.013737 | 0.014752 | 0.014401 | 0.013462 | 0.013177 | 0.014027 | 0.013502 | 0.013801 | 0.00054989 | 3.9843 | 0.0031901 | 30.0635
Gross counts in the ROI: | [17 empty] | Gross counts in the ROI, spectrum manually recalibrated using the 0 and Cu peaks
| Analysis | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Average | Stdev | % error |  | Spectrum | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |  | Average | Stdev |  | % error
| Na | 22871 | 22801 | 23107 | 22846 | 22995 | 22781 | 22972 | 22971 | 22851 | 23062 | 22925.7 | 111.484 | 0.48628 |  | Na | 20159 | 20130 | 20317 | 20137 | 20306 | 20066 | 20227 | 20101 | 20083 | 20292 |  | 20181.8 | 95.9176 |  | 0.47527
| Mg | 27202 | 27483 | 27061 | 27259 | 27257 | 27316 | 27146 | 27656 | 27488 | 27380 | 27324.8 | 179.5196 | 0.65698 |  | Mg | 22580 | 24481 | 24026 | 22597 | 24170 | 24315 | 24084 | 22894 | 24457 | 24339 |  | 23794.3 | 780.0219 |  | 3.2782
| Al | 234560 | 237466 | 238618 | 237956 | 238276 | 237537 | 238704 | 237529 | 236278 | 236588 | 237351.2 | 1258.7574 | 0.53034 |  | Al | 221989 | 224669 | 225755 | 225230 | 225501 | 224775 | 225906 | 224815 | 223547 | 223781 |  | 224596.8 | 1198.4602 |  | 0.53361
| Si | 1632563 | 1649776 | 1649641 | 1654771 | 1650485 | 1647758 | 1652665 | 1652609 | 1651429 | 1646277 | 1648797.4 | 6214.6928 | 0.37692 |  | Si | 1476455 | 1584433 | 1538116 | 1496320 | 1538826 | 1490254 | 1541415 | 1541729 | 1492747 | 1488580 |  | 1518887.5 | 34679.9184 |  | 2.2832
| P | 25171 | 25104 | 25270 | 25375 | 25080 | 25497 | 25157 | 25435 | 25095 | 25273 | 25245.7 | 149.3445 | 0.59156 |  | P | 19886 | 22512 | 22537 | 20100 | 22402 | 21385 | 22534 | 21549 | 19800 | 20060 |  | 21276.5 | 1203.0966 |  | 5.6546
| K | 831505 | 836381 | 837651 | 838006 | 835946 | 837890 | 837269 | 837928 | 836630 | 836412 | 836561.8 | 1926.8003 | 0.23032 |  | K | 732455 | 784688 | 765487 | 738176 | 763682 | 738447 | 765304 | 737566 | 735904 | 736256 |  | 749796.5 | 18219.0524 |  | 2.4299
| Ca | 435435 | 437862 | 438519 | 440867 | 439239 | 438718 | 438771 | 439303 | 438873 | 437641 | 438522.8 | 1397.936 | 0.31878 |  | Ca | 360789 | 394306 | 385135 | 387541 | 385951 | 364306 | 385601 | 386172 | 385305 | 362689 |  | 379779.5 | 12181.3493 |  | 3.2075
| Ti | 227066 | 229033 | 228950 | 229157 | 230391 | 229766 | 230777 | 230460 | 230179 | 229304 | 229508.3 | 1079.0082 | 0.47014 |  | Ti | 192620 | 200408 | 199962 | 194545 | 195627 | 194942 | 195784 | 195361 | 195104 | 194465 |  | 195881.8 | 2435.8921 |  | 1.2436
| V | 84477 | 85947 | 85657 | 85458 | 85341 | 85228 | 85391 | 85971 | 85277 | 85340 | 85408.7 | 421.3352 | 0.49332 |  | V | 60038 | 63662 | 63500 | 60475 | 63046 | 63169 | 63361 | 60994 | 60487 | 60617 |  | 61934.9 | 1515.5426 |  | 2.447
| Cr | 50601 | 51342 | 51212 | 51133 | 51134 | 51058 | 51435 | 51322 | 50778 | 51055 | 51107 | 256.2217 | 0.50134 |  | Cr | 34174 | 37105 | 36931 | 34378 | 36954 | 36765 | 36972 | 37037 | 36682 | 34529 |  | 36152.7 | 1245.6249 |  | 3.4455
| Mn | 95126 | 96401 | 96409 | 96272 | 96681 | 96171 | 96454 | 96739 | 96855 | 96555 | 96366.3 | 484.0195 | 0.50227 |  | Mn | 67702 | 76107 | 76160 | 68575 | 72491 | 71925 | 72361 | 68749 | 68557 | 68851 |  | 71147.8 | 3159.2986 |  | 4.4405
| Fe | 4762965 | 4807648 | 4810355 | 4815815 | 4810380 | 4811755 | 4823348 | 4826131 | 4812440 | 4810273 | 4809111 | 17293.1939 | 0.35959 |  | Fe | 3860185 | 4288560 | 4184901 | 3901971 | 4029374 | 4031600 | 4037672 | 4038522 | 4026884 | 4022758 |  | 4042242.7 | 122371.8865 |  | 3.0273
| Co | 402615 | 407277 | 407998 | 408481 | 406514 | 407749 | 409921 | 411280 | 411390 | 410995 | 408422 | 2690.5894 | 0.65878 |  | Co | 245717 | 296025 | 296822 | 254044 | 295931 | 292484 | 298384 | 255849 | 296057 | 295317 |  | 282663 | 21450.7955 |  | 7.5888
| Ni | 58203 | 58895 | 59556 | 59152 | 59253 | 59081 | 58979 | 58676 | 58625 | 58901 | 58932.1 | 373.985 | 0.6346 |  | Ni | 37839 | 43134 | 41226 | 38619 | 41001 | 40949 | 40907 | 38309 | 40514 | 38459 |  | 40095.7 | 1700.7886 |  | 4.2418
| Cu | 74795 | 75742 | 75926 | 75991 | 75485 | 75973 | 76028 | 75710 | 75250 | 75558 | 75645.8 | 391.3327 | 0.51732 |  | Cu | 49327 | 55847 | 55942 | 52925 | 52701 | 53163 | 53015 | 50035 | 49744 | 49775 |  | 52247.4 | 2460.868 |  | 4.71
| Zn | 106488 | 107597 | 107521 | 108371 | 107036 | 108023 | 108337 | 107944 | 108114 | 106634 | 107606.5 | 682.5357 | 0.63429 |  | Zn | 70011 | 78545 | 78499 | 75084 | 74296 | 75090 | 78846 | 74580 | 74840 | 70197 |  | 74998.8 | 3131.798 |  | 4.1758
| Ga | 95653 | 96576 | 96430 | 96546 | 96825 | 96475 | 96331 | 96442 | 96333 | 96416 | 96402.7 | 299.7495 | 0.31093 |  | Ga | 59596 | 71274 | 67495 | 63675 | 68103 | 63922 | 67575 | 63688 | 63728 | 63797 |  | 65285.3 | 3298.8354 |  | 5.053
| Rb | 108935 | 110870 | 111096 | 111383 | 111082 | 110790 | 111094 | 111118 | 111203 | 111098 | 110866.9 | 697.955 | 0.62954 |  | Rb | 69434 | 80568 | 77562 | 67813 | 77476 | 73903 | 77404 | 70910 | 74036 | 70061 |  | 73916.7 | 4263.9148 |  | 5.7685
| Sr | 95953 | 96677 | 97236 | 97087 | 97060 | 96675 | 97099 | 97448 | 97868 | 97153 | 97025.6 | 511.4046 | 0.52708 |  | Sr | 60415 | 70011 | 70399 | 61731 | 64556 | 63676 | 67363 | 61367 | 64521 | 60797 |  | 64483.6 | 3680.707 |  | 5.708
| Y | 62238 | 62693 | 62711 | 63101 | 62865 | 63001 | 62974 | 63166 | 62976 | 62818 | 62854.3 | 265.934 | 0.4231 |  | Y | 36462 | 44295 | 42485 | 38828 | 40636 | 38623 | 40576 | 38729 | 38732 | 36582 |  | 39594.8 | 2458.7666 |  | 6.2098
| Zr | 107967 | 108876 | 109425 | 109955 | 109816 | 110944 | 111703 | 111521 | 112277 | 112340 | 110482.4 | 1498.1837 | 1.356 |  | Zr | 68088 | 84317 | 78732 | 73040 | 79050 | 73018 | 77175 | 73967 | 73947 | 70148 |  | 75148.2 | 4748.231 |  | 6.3185
| Nb | 39380 | 39747 | 39822 | 40132 | 40306 | 39955 | 39569 | 39984 | 40174 | 39654 | 39872.3 | 291.7598 | 0.73174 |  | Nb | 21751 | 27674 | 25474 | 23313 | 25857 | 24435 | 25448 | 24290 | 23537 | 22008 |  | 24378.7 | 1818.6593 |  | 7.46
| Ba | 211792 | 213514 | 213674 | 213608 | 214717 | 214304 | 215646 | 215568 | 215822 | 214823 | 214346.8 | 1244.5506 | 0.58062 |  | Ba | 179847 | 184794 | 184953 | 168042 | 185631 | 182150 | 183364 | 183223 | 183623 | 182605 |  | 181823.2 | 5112.4931 |  | 2.8118
| La | 90280 | 91187 | 90743 | 91148 | 92006 | 91330 | 91080 | 90805 | 89832 | 89930 | 90834.1 | 670.0397 | 0.73765 |  | La | 42924 | 53881 | 51849 | 43420 | 52551 | 43565 | 43176 | 43296 | 42644 | 42572 |  | 45987.8 | 4709.1281 |  | 10.24
| Ce | 79764 | 80769 | 81064 | 80355 | 80982 | 80897 | 80724 | 81411 | 81366 | 81254 | 80858.6 | 500.6374 | 0.61915 |  | Ce | 58137 | 65903 | 58835 | 58434 | 59007 | 58748 | 58921 | 59339 | 59268 | 59162 |  | 59575.4 | 2253.7793 |  | 3.7831
| Th | 82245 | 83413 | 83233 | 82681 | 84194 | 83056 | 83293 | 82953 | 83410 | 82961 | 83143.9 | 513.8118 | 0.61798 |  | Th | 46775 | 55858 | 55238 | 46835 | 53457 | 52604 | 55593 | 49611 | 52627 | 49564 |  | 51816.2 | 3439.0033 |  | 6.6369
Net counts in the ROI (deconvoluted peaks):
| Analysis | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Average | Stdev |  | % error
| Na | 3052 | 2987 | 3332 | 2963 | 3158 | 2959 | 3176 | 2920 | 2892 | 3209 | 3064.8 | 145.9991 |  | 4.7637
| Mg | 2229 | 2225 | 2199 | 2172 | 2091 | 2396 | 2184 | 2371 | 2214 | 2073 | 2215.4 | 103.2787 |  | 4.6619
| Al | 199982 | 202674 | 204238 | 203446 | 203594 | 202993 | 204476 | 203264 | 201639 | 201925 | 202823.1 | 1343.3923 |  | 0.66235
| Si | 1685077 | 1702677 | 1703288 | 1708000 | 1704204 | 1700877 | 1706192 | 1705312 | 1704016 | 1698918 | 1701856.1 | 6433.6102 |  | 0.37803
| P | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| K | 984777 | 989230 | 992006 | 991651 | 989515 | 991208 | 995265 | 997834 | 997111 | 997190 | 992578.7 | 4231.8136 |  | 0.42635
| Ca | 426658 | 430119 | 429983 | 433351 | 430670 | 430613 | 430227 | 431711 | 430490 | 429275 | 430309.7 | 1700.4221 |  | 0.39516
| Ti | 215761 | 218685 | 216818 | 217635 | 219017 | 217438 | 215808 | 214918 | 213459 | 212875 | 216241.4 | 2071.6602 |  | 0.95803
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Cr | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Mn | 32938 | 32987 | 33807 | 34057 | 34276 | 33019 | 33690 | 33201 | 33512 | 33842 | 33532.9 | 477.1248 |  | 1.4229
| Fe | 5733305 | 5790509 | 5794212 | 5798642 | 5792444 | 5792847 | 5801312 | 5802766 | 5787901 | 5784131 | 5787806.9 | 20007.6984 |  | 0.34569
| Co | 35608 | 35506 | 35134 | 36226 | 34897 | 35631 | 35927 | 37324 | 38364 | 36691 | 36130.8 | 1064.8264 |  | 2.9471
| Ni | 1920 | 2585 | 2409 | 2628 | 1870 | 1817 | 2245 | 1785 | 2071 | 1565 | 2089.5 | 362.2603 |  | 17.3372
| Cu | 11390 | 11600 | 11369 | 11888 | 11479 | 11839 | 11967 | 12127 | 11432 | 11674 | 11676.5 | 266.6172 |  | 2.2834
| Zn | 32756 | 33265 | 32409 | 33525 | 32530 | 33368 | 33866 | 33466 | 33931 | 32137 | 33125.3 | 626.5099 |  | 1.8913
| Ga | 7279 | 7410 | 7028 | 6788 | 7933 | 7409 | 7131 | 7525 | 7036 | 7145 | 7268.4 | 320.4442 |  | 4.4087
| Rb | 45726 | 47055 | 47093 | 47851 | 47314 | 47652 | 47011 | 47450 | 47522 | 47767 | 47244.1 | 610.1181 |  | 1.2914
| Sr | 43255 | 43264 | 44067 | 43765 | 43870 | 43544 | 43741 | 44194 | 44744 | 44587 | 43903.1 | 504.4466 |  | 1.149
| Y | 3318 | 3373 | 3235 | 3470 | 2767 | 3242 | 2855 | 3399 | 3183 | 2979 | 3182.1 | 238.4633 |  | 7.4939
| Zr | 72594 | 73437 | 73530 | 74527 | 74070 | 75706 | 76460 | 75949 | 77030 | 77600 | 75090.3 | 1693.4951 |  | 2.2553
| Nb | 6064 | 5178 | 5893 | 6209 | 6064 | 5920 | 5550 | 5524 | 6066 | 5463 | 5793.1 | 339.7926 |  | 5.8655
| Ba | 66113 | 69111 | 68811 | 70689 | 68873 | 74281 | 79961 | 80731 | 80466 | 77676 | 73671.2 | 5631.6718 |  | 7.6443
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Ce | 10793 | 9164 | 11133 | 10181 | 9961 | 8942 | 5628 | 6775 | 8945 | 9854 | 9137.6 | 1731.3159 |  | 18.9472
| Th | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Table S5 AGV-2
Non-calibrated data (wt%): AGV-2
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) | Calibration Factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 4.8304 | 4.5676 | 4.6897 | 4.4552 | 4.4747 | 4.4568 | 4.3172 | 4.1264 | 4.6913 | 4.2484 | 4.4858 | 0.21733 | 4.8449 | 0.24691 | 5.8249 | 0.94496
| MgO | 0.64979 | 0.6538 | 0.58324 | 0.66618 | 0.66047 | 0.55105 | 0.5661 | 0.63079 | 0.64431 | 0.54639 | 0.61521 | 0.047967 | 7.7969 | -1.1957 | -66.0269 | 2.9435
| Al2O3 | 16.5948 | 16.6307 | 16.6103 | 16.6737 | 16.5605 | 16.5824 | 16.6729 | 16.6723 | 16.5429 | 16.6436 | 16.6184 | 0.047952 | 0.28855 | -0.48884 | -2.8575 | 1.0294
| SiO2 | 60.1246 | 60.2697 | 60.2456 | 60.3002 | 60.3602 | 60.4209 | 60.3861 | 60.5006 | 60.2058 | 60.4773 | 60.3291 | 0.12128 | 0.20104 | 0.33739 | 0.56239 | 0.99441
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.4856 | -100 | nd
| K2O | 3.8972 | 3.9119 | 3.9118 | 3.9173 | 3.921 | 3.926 | 3.9209 | 3.9345 | 3.9036 | 3.9519 | 3.9196 | 0.015612 | 0.3983 | 1.006 | 34.5287 | 0.74334
| CaO | 5.0113 | 5.0427 | 5.0345 | 5.0376 | 5.0597 | 5.0666 | 5.0681 | 5.0794 | 5.0377 | 5.0688 | 5.0506 | 0.021096 | 0.4177 | -0.21001 | -3.9921 | 1.0416
| TiO2 | 1.0695 | 1.071 | 1.0738 | 1.0706 | 1.0748 | 1.0828 | 1.0911 | 1.0878 | 1.0726 | 1.0932 | 1.0787 | 0.0091335 | 0.8467 | 0.016463 | 1.5499 | 0.98474
| V | 0.0079572 | 0.009113 | 0.0091885 | 0.0098024 | 0.008333 | 0.0071907 | 0.0070578 | 0.0091652 | 0.010221 | 0.010166 | 0.0088195 | 0.0011461 | 12.9951 | -0.0033205 | -27.3518 | 1.3765
| Cr | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0017198 | -100 | nd
| MnO | 0.11214 | 0.11444 | 0.11277 | 0.11341 | 0.11365 | 0.1141 | 0.1148 | 0.11542 | 0.11444 | 0.11442 | 0.11396 | 0.00098023 | 0.86016 | 0.013382 | 13.3055 | 0.88257
| Fe2O3t | 7.4644 | 7.5036 | 7.5009 | 7.5202 | 7.553 | 7.5727 | 7.599 | 7.6096 | 7.5278 | 7.6154 | 7.5467 | 0.05156 | 0.68322 | 0.77864 | 11.5046 | 0.89682
| Co | 0.025069 | 0.024821 | 0.025079 | 0.025365 | 0.025963 | 0.025085 | 0.025829 | 0.026222 | 0.025655 | 0.025146 | 0.025424 | 0.00046522 | 1.8299 | 0.023805 | 1470.6497 | 0.063668
| Ni | 0.0013506 | 0.0011424 | 0.0012542 | 0.0008911 | 0.0012669 | 0.0011521 | 0.00083611 | 0.00092893 | 0.00098627 | 0.0012504 | 0.0011059 | 0.00018156 | 16.4174 | -0.00081628 | -42.4666 | 1.7381
| Cu | 0.007571 | 0.0074274 | 0.0078007 | 0.0076682 | 0.007328 | 0.0073894 | 0.0079458 | 0.0074961 | 0.0076698 | 0.0076206 | 0.0075917 | 0.00019116 | 2.518 | 0.0022299 | 41.5882 | 0.70627
| Zn | 0.009838 | 0.0098598 | 0.01019 | 0.0099136 | 0.0098821 | 0.010362 | 0.010222 | 0.010289 | 0.010416 | 0.0099104 | 0.010088 | 0.00022863 | 2.2663 | 0.001388 | 15.9539 | 0.86241
| Ga | 0.0032353 | 0.003062 | 0.0035883 | 0.0034462 | 0.0037183 | 0.0037381 | 0.0032363 | 0.0034768 | 0.0032571 | 0.0033424 | 0.0034101 | 0.00022362 | 6.5575 | 0.0013867 | 68.5379 | 0.59334
| Rb | 0.0096025 | 0.0099558 | 0.010135 | 0.0098312 | 0.0094281 | 0.010296 | 0.010138 | 0.0099882 | 0.010566 | 0.010498 | 0.010044 | 0.00036357 | 3.6198 | 0.0031038 | 44.7239 | 0.69097
| Sr | 0.081173 | 0.081004 | 0.080466 | 0.081203 | 0.081179 | 0.082279 | 0.08261 | 0.082535 | 0.08103 | 0.082147 | 0.081563 | 0.00075459 | 0.92517 | 0.014995 | 22.5261 | 0.81615
| Y | 0.0015865 | 0.0016568 | 0.0015356 | 0.00091316 | 0.0015116 | 0.0010869 | 0.0016781 | 0.0015623 | 0.0016334 | 0.0011009 | 0.0014265 | 0.00028031 | 19.6499 | -0.0005968 | -29.4959 | 1.4184
| Zr | 0.035096 | 0.035313 | 0.035845 | 0.035489 | 0.036471 | 0.035606 | 0.035729 | 0.036456 | 0.035957 | 0.036714 | 0.035868 | 0.00053423 | 1.4894 | 0.012599 | 54.148 | 0.64873
| Nb | 0.0040615 | 0.0037446 | 0.0033619 | 0.0036037 | 0.0039594 | 0.0037248 | 0.0035646 | 0.0039304 | 0.0039456 | 0.0041128 | 0.0038009 | 0.00024138 | 6.3505 | 0.0022834 | 150.4742 | 0.39924
| Ba | 0.1983 | 0.19941 | 0.20114 | 0.19902 | 0.1821 | 0.19665 | 0.20761 | 0.20617 | 0.19166 | 0.20333 | 0.19854 | 0.007398 | 3.7262 | 0.08321 | 72.1494 | 0.58089
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0038443 | -100 | nd
| Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0068793 | -100 | nd
| Th | 0.011205 | 0.012 | 0.011047 | 0.011463 | 0.010533 | 0.011326 | 0.010443 | 0.010711 | 0.011885 | 0.011279 | 0.011189 | 0.00052496 | 4.6916 | 0.010572 | 1713.1578 | 0.055152
Calibrated data (wt%): AGV-2
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 3.9779 | 4.3814 | 3.8526 | 4.0409 | 4.1725 | 4.1742 | 4.1557 | 4.3871 | 4.2626 | 4.51 | 4.1915 | 0.20164 | 4.8107 | -0.04739 | -1.118
| MgO | 1.5942 | 1.8747 | 1.8352 | 1.651 | 1.6075 | 1.9197 | 1.9361 | 1.6998 | 1.901 | 1.89 | 1.7909 | 0.13695 | 7.6468 | -0.019968 | -1.1027
| Al2O3 | 17.2116 | 17.094 | 17.2227 | 17.2415 | 17.1461 | 17.1012 | 17.218 | 17.1706 | 17.177 | 17.1435 | 17.1726 | 0.05131 | 0.29879 | 0.065358 | 0.38205
| SiO2 | 60.4857 | 60.1334 | 60.4179 | 60.3867 | 60.416 | 60.2413 | 60.1828 | 60.2183 | 60.1704 | 60.0365 | 60.2689 | 0.14815 | 0.24581 | 0.27717 | 0.46202
| P2O5 | 0.54206 | 0.5339 | 0.53726 | 0.53848 | 0.53921 | 0.53637 | 0.5345 | 0.5356 | 0.53458 | 0.53145 | 0.53634 | 0.0030409 | 0.56697 | 0.050741 | 10.4491
| K2O | 2.9552 | 2.9123 | 2.935 | 2.9312 | 2.9351 | 2.922 | 2.9191 | 2.9225 | 2.9167 | 2.9065 | 2.9256 | 0.014035 | 0.47972 | 0.011976 | 0.41105
| CaO | 5.2233 | 5.1795 | 5.2208 | 5.2216 | 5.2204 | 5.1959 | 5.1729 | 5.1836 | 5.1808 | 5.1505 | 5.1949 | 0.02552 | 0.49125 | -0.065724 | -1.2493
| TiO2 | 1.0681 | 1.0466 | 1.059 | 1.0645 | 1.0582 | 1.0487 | 1.0464 | 1.0486 | 1.0404 | 1.0434 | 1.0524 | 0.0093989 | 0.89311 | -0.009868 | -0.92897
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0.0010497 | 0 | 0 | 0 | 0.00010497 | 0.00033195 | 316.2278 | -0.012035 | -99.1353
| Cr | 0.0028754 | 0.0030432 | 0.0031737 | 0.0032454 | 0.0024893 | 0.0031698 | 0.0030222 | 0.0028594 | 0.0025472 | 0.00328 | 0.0029706 | 0.00027737 | 9.3374 | 0.0012507 | 72.7247
| MnO | 0.095896 | 0.095664 | 0.096698 | 0.096754 | 0.09579 | 0.095164 | 0.094898 | 0.094561 | 0.095771 | 0.093855 | 0.095505 | 0.00090744 | 0.95014 | -0.0050714 | -5.0423
| Fe2O3t | 6.6538 | 6.5609 | 6.6283 | 6.6366 | 6.6182 | 6.5781 | 6.5462 | 6.5483 | 6.5331 | 6.5037 | 6.5807 | 0.05051 | 0.76754 | -0.18733 | -2.7678
| Co | 0.0014459 | 0.0014717 | 0.0015034 | 0.0014853 | 0.0014431 | 0.0014886 | 0.0014531 | 0.0014411 | 0.001422 | 0.0014377 | 0.0014592 | 2.6461e-05 | 1.8134 | -0.00015949 | -9.8529
| Ni | 0.0020687 | 0.0016255 | 0.0015293 | 0.0013804 | 0.0019061 | 0.0020872 | 0.0014661 | 0.0020732 | 0.0018822 | 0.0022279 | 0.0018246 | 0.00030104 | 16.4984 | -9.7516e-05 | -5.0732
| Cu | 0.0051487 | 0.0051672 | 0.0050468 | 0.0053657 | 0.004994 | 0.0049326 | 0.0051591 | 0.005266 | 0.0049988 | 0.0050996 | 0.0051179 | 0.00013236 | 2.5863 | -0.00024396 | -4.55
| Zn | 0.0081026 | 0.0084914 | 0.0083825 | 0.0083533 | 0.0084744 | 0.0080494 | 0.0080711 | 0.0083243 | 0.0080299 | 0.0080187 | 0.0082298 | 0.00019269 | 2.3414 | -0.00047056 | -5.4085
| Ga | 0.001873 | 0.0018199 | 0.0019414 | 0.0018126 | 0.0020954 | 0.0020758 | 0.001923 | 0.0020091 | 0.0017091 | 0.0018074 | 0.0019067 | 0.00012587 | 6.6015 | -0.00011668 | -5.7667
| Rb | 0.0068415 | 0.0068659 | 0.0064855 | 0.0066035 | 0.0067122 | 0.0061215 | 0.0063794 | 0.0065994 | 0.0064624 | 0.0062386 | 0.006531 | 0.00024365 | 3.7307 | -0.00040905 | -5.894
| Sr | 0.063515 | 0.062466 | 0.06358 | 0.063837 | 0.06364 | 0.062532 | 0.062513 | 0.062163 | 0.062381 | 0.062566 | 0.062919 | 0.00063786 | 1.0138 | -0.0036484 | -5.4807
| Y | 0.0014666 | 0.0021723 | 0.0020761 | 0.0022373 | 0.0014486 | 0.0020087 | 0.0012103 | 0.0020465 | 0.0022014 | 0.0021094 | 0.0018977 | 0.0003733 | 19.6708 | -0.0001256 | -6.2074
| Zr | 0.022439 | 0.021911 | 0.022199 | 0.021825 | 0.02177 | 0.022208 | 0.021596 | 0.021891 | 0.021497 | 0.021383 | 0.021872 | 0.00033526 | 1.5328 | -0.0013964 | -6.0013
| Nb | 0.0015474 | 0.00148 | 0.0014733 | 0.0013405 | 0.0014021 | 0.0014841 | 0.0013501 | 0.0012641 | 0.0014032 | 0.0015231 | 0.0014268 | 9.008e-05 | 6.3135 | -9.0713e-05 | -5.9778
| Ba | 0.064266 | 0.059953 | 0.06699 | 0.063563 | 0.064003 | 0.061395 | 0.067469 | 0.06468 | 0.061929 | 0.064421 | 0.063867 | 0.002337 | 3.6592 | -0.051463 | -44.6224
| Th | 0.010645 | 0.011184 | 0.010075 | 0.0098544 | 0.010693 | 0.0099075 | 0.010774 | 0.010418 | 0.01128 | 0.010545 | 0.010538 | 0.00048909 | 4.6414 | 0.0099205 | 1607.5566

## Table S6 BHVO-2
Non-calibrated data (wt%): BHVO-2
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) | Calibration factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 2.3277 | 2.2158 | 1.9053 | 2.1328 | 2.043 | 2.0282 | 1.8351 | 2.2569 | 2.1028 | 1.9684 | 2.0816 | 0.15682 | 7.5338 | -0.12916 | -5.8423 | 1.062
| MgO | 4.9584 | 5.0208 | 5.0075 | 5.0257 | 5.1312 | 4.9995 | 5.0413 | 4.9615 | 5.1108 | 5.0083 | 5.0265 | 0.056385 | 1.1218 | -2.1734 | -30.1866 | 1.4324
| Al2O3 | 15.0492 | 15.112 | 15.198 | 15.0949 | 15.1474 | 15.1642 | 15.1949 | 15.1216 | 15.1845 | 15.122 | 15.1389 | 0.047958 | 0.31679 | 1.695 | 12.6084 | 0.88803
| SiO2 | 48.6301 | 48.5768 | 48.6991 | 48.6854 | 48.6814 | 48.8131 | 48.8216 | 48.6387 | 48.6179 | 48.7558 | 48.692 | 0.082434 | 0.1693 | -1.0004 | -2.0131 | 1.0205
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.26888 | -100 | nd
| K2O | 0.69889 | 0.70324 | 0.70637 | 0.6948 | 0.6931 | 0.69521 | 0.69532 | 0.69156 | 0.69353 | 0.69982 | 0.69719 | 0.0047809 | 0.68574 | 0.17935 | 34.6343 | 0.74275
| CaO | 11.2335 | 11.2365 | 11.2619 | 11.2297 | 11.2196 | 11.2202 | 11.2622 | 11.2197 | 11.2032 | 11.2767 | 11.2363 | 0.023371 | 0.20799 | -0.11623 | -1.0238 | 1.0103
| TiO2 | 2.797 | 2.8106 | 2.8302 | 2.8041 | 2.8102 | 2.8087 | 2.816 | 2.8162 | 2.809 | 2.8261 | 2.8128 | 0.0098202 | 0.34912 | 0.094169 | 3.4638 | 0.96652
| V | 0.035436 | 0.032246 | 0.034126 | 0.032618 | 0.036997 | 0.032513 | 0.036765 | 0.03399 | 0.035334 | 0.035022 | 0.034505 | 0.0017056 | 4.9432 | 0.0029367 | 9.3027 | 0.91489
| Cr | 0.029426 | 0.028691 | 0.030068 | 0.028994 | 0.028337 | 0.030214 | 0.028745 | 0.029026 | 0.028623 | 0.029699 | 0.029182 | 0.0006407 | 2.1955 | 0.0012989 | 4.6582 | 0.95549
| MnO | 0.18058 | 0.17921 | 0.18178 | 0.18043 | 0.17957 | 0.17989 | 0.18111 | 0.18021 | 0.18059 | 0.18086 | 0.18042 | 0.00075188 | 0.41674 | 0.01456 | 8.7781 | 0.9193
| Fe2O3t | 13.9251 | 13.9288 | 13.9997 | 13.9408 | 13.8949 | 13.8815 | 13.9587 | 13.8885 | 13.8862 | 13.9826 | 13.9287 | 0.041895 | 0.30078 | 1.6799 | 13.7146 | 0.87939
| Co | 0.044854 | 0.045868 | 0.044842 | 0.044525 | 0.044791 | 0.044886 | 0.045645 | 0.045186 | 0.043687 | 0.044957 | 0.044924 | 0.0005963 | 1.3273 | 0.040443 | 902.4865 | 0.099752
| Ni | 0.0099844 | 0.010154 | 0.010661 | 0.010041 | 0.010326 | 0.010151 | 0.010429 | 0.0097189 | 0.010331 | 0.010606 | 0.01024 | 0.0002894 | 2.8262 | -0.0016103 | -13.5885 | 1.1573
| Cu | 0.016402 | 0.016693 | 0.017143 | 0.016591 | 0.016403 | 0.017158 | 0.016834 | 0.016158 | 0.016549 | 0.016627 | 0.016656 | 0.00031883 | 1.9143 | 0.0040085 | 31.6947 | 0.75933
| Zn | 0.011831 | 0.011941 | 0.011869 | 0.012241 | 0.011679 | 0.012037 | 0.0117 | 0.01114 | 0.01148 | 0.011898 | 0.011782 | 0.00030636 | 2.6003 | 0.0015245 | 14.8626 | 0.87061
| Ga | 0.0031656 | 0.0035598 | 0.0033229 | 0.0032689 | 0.0029877 | 0.0034279 | 0.0037343 | 0.0033804 | 0.0033688 | 0.003816 | 0.0034032 | 0.00024955 | 7.3326 | 0.0012423 | 57.4865 | 0.63498
| Rb | 0.008255 | 0.0092144 | 0.0084255 | 0.0085482 | 0.0086298 | 0.0082587 | 0.0085552 | 0.0081974 | 0.0083574 | 0.0081764 | 0.0084618 | 0.00030884 | 3.6498 | 0.0074859 | 767.0559 | 0.11533
| Sr | 0.047341 | 0.048053 | 0.047595 | 0.047316 | 0.046244 | 0.04547 | 0.046889 | 0.046886 | 0.047499 | 0.047615 | 0.047091 | 0.00075735 | 1.6083 | 0.0083527 | 21.562 | 0.82263
| Y | 0.0013918 | 0.0011272 | 0.0016993 | 0.0012841 | 0.00086556 | 0.0014951 | 0.0012917 | 0.0016208 | 0.00095119 | 0.0014489 | 0.0013176 | 0.00027254 | 20.6848 | -0.0012716 | -49.1127 | 1.9651
| Zr | 0.02985 | 0.028724 | 0.030534 | 0.029347 | 0.028279 | 0.028163 | 0.029867 | 0.029179 | 0.028532 | 0.028891 | 0.029137 | 0.00076893 | 2.639 | 0.012008 | 70.1068 | 0.58787
| Nb | 0.0043546 | 0.0045615 | 0.004931 | 0.004432 | 0.0046131 | 0.0047518 | 0.0053823 | 0.0044745 | 0.0046476 | 0.004541 | 0.0046689 | 0.00029975 | 6.42 | 0.0028764 | 160.4693 | 0.38392
| Ba | 0.11636 | 0.14055 | 0.1199 | 0.13192 | 0.12714 | 0.13496 | 0.10998 | 0.13782 | 0.1245 | 0.10886 | 0.1252 | 0.011275 | 9.0058 | 0.11225 | 867.0917 | 0.1034
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0014938 | -100 | nd
| Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0037842 | -100 | nd
| Th | 0.019051 | 0.019263 | 0.019044 | 0.0197 | 0.018845 | 0.019663 | 0.019341 | 0.019527 | 0.019487 | 0.019515 | 0.019344 | 0.00028723 | 1.4849 | 0.019224 | 16087.021 | 0.0061778
Calibrated Data (wt%): BHVO-2
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 1.6292 | 1.7368 | 1.8718 | 1.5173 | 1.6795 | 1.6864 | 1.7656 | 1.5772 | 1.835 | 1.9303 | 1.7229 | 0.13136 | 7.6242 | -0.48786 | -22.0676
| MgO | 12.9447 | 13.1815 | 12.8448 | 13.0192 | 12.9271 | 13.2259 | 12.9888 | 12.9456 | 12.98 | 12.8354 | 12.9893 | 0.12747 | 0.98138 | 5.7894 | 80.4091
| Al2O3 | 14.8463 | 14.8951 | 14.8603 | 14.9126 | 14.8921 | 14.8539 | 14.8222 | 14.9252 | 14.842 | 14.7867 | 14.8636 | 0.042813 | 0.28804 | 1.4198 | 10.5611
| SiO2 | 45.644 | 45.4648 | 45.5869 | 45.6813 | 45.712 | 45.5084 | 45.5874 | 45.6023 | 45.4984 | 45.5715 | 45.5857 | 0.079725 | 0.17489 | -4.1066 | -8.2641
| P2O5 | 0.070854 | 0.070322 | 0.070477 | 0.070628 | 0.070355 | 0.070353 | 0.07101 | 0.071039 | 0.070943 | 0.071081 | 0.070706 | 0.00031212 | 0.44143 | -0.19817 | -73.7031
| K2O | 0.47456 | 0.46952 | 0.46979 | 0.47111 | 0.47165 | 0.46896 | 0.47122 | 0.47913 | 0.47719 | 0.47466 | 0.47278 | 0.0034509 | 0.72991 | -0.045055 | -8.7006
| CaO | 10.6751 | 10.5877 | 10.6397 | 10.6523 | 10.626 | 10.5975 | 10.633 | 10.663 | 10.643 | 10.6512 | 10.6369 | 0.02734 | 0.25703 | -0.7157 | -6.3043
| TiO2 | 2.5143 | 2.4925 | 2.5043 | 2.5013 | 2.4975 | 2.4934 | 2.4989 | 2.5145 | 2.5006 | 2.4919 | 2.5009 | 0.0081755 | 0.3269 | -0.21773 | -8.0089
| V | 0.024622 | 0.026056 | 0.026556 | 0.026067 | 0.023557 | 0.026598 | 0.021798 | 0.025314 | 0.022645 | 0.02678 | 0.024999 | 0.0017801 | 7.1207 | -0.0065687 | -20.8081
| Cr | 0.032614 | 0.031705 | 0.031912 | 0.03191 | 0.032664 | 0.031936 | 0.032474 | 0.032593 | 0.032106 | 0.031692 | 0.032161 | 0.00038712 | 1.2037 | 0.0042771 | 15.3393
| MnO | 0.13571 | 0.13519 | 0.13569 | 0.13559 | 0.13526 | 0.13409 | 0.13527 | 0.13658 | 0.13451 | 0.13602 | 0.13539 | 0.00071403 | 0.52738 | -0.030471 | -18.3712
| Fe2O3 | 10.8934 | 10.7909 | 10.8367 | 10.8632 | 10.8131 | 10.7884 | 10.8565 | 10.9063 | 10.8479 | 10.872 | 10.8468 | 0.040106 | 0.36975 | -1.402 | -11.4458
| Co | 0.0022622 | 0.0021923 | 0.0022787 | 0.0022959 | 0.0022604 | 0.0022459 | 0.0022403 | 0.0022568 | 0.0023094 | 0.0022637 | 0.0022606 | 3.2112e-05 | 1.4205 | -0.0022207 | -49.5556
| Ni | 0.015208 | 0.014773 | 0.013961 | 0.014933 | 0.014556 | 0.01475 | 0.014393 | 0.015286 | 0.014555 | 0.014351 | 0.014677 | 0.00040286 | 2.7449 | 0.0028261 | 23.8478
| Cu | 0.0097167 | 0.0096447 | 0.009462 | 0.0098253 | 0.010029 | 0.00955 | 0.0096932 | 0.010018 | 0.0097544 | 0.0096101 | 0.0097303 | 0.00018539 | 1.9052 | -0.0029168 | -23.063
| Zn | 0.0084085 | 0.0080903 | 0.0078887 | 0.008258 | 0.0085079 | 0.0082221 | 0.0086475 | 0.0083877 | 0.0084374 | 0.008382 | 0.008323 | 0.00021722 | 2.6099 | -0.0019341 | -18.8566
| Ga | 0.0018469 | 0.0016259 | 0.0016395 | 0.0018051 | 0.0016593 | 0.0014406 | 0.0015816 | 0.0016083 | 0.0017228 | 0.0015362 | 0.0016466 | 0.00012111 | 7.3549 | -0.00051434 | -23.8011
| Rb | 0.0045863 | 0.0046741 | 0.0046075 | 0.0047925 | 0.0046328 | 0.0048214 | 0.0047929 | 0.0047259 | 0.0051671 | 0.0046427 | 0.0047443 | 0.00017032 | 3.5901 | 0.0037684 | 386.1372
| Sr | 0.031683 | 0.031512 | 0.031261 | 0.031158 | 0.030258 | 0.030649 | 0.031471 | 0.031668 | 0.031966 | 0.031584 | 0.031321 | 0.00051723 | 1.6514 | -0.0074172 | -19.1471
| Y | 0.0016616 | 0.0010855 | 0.0018634 | 0.0014784 | 0.0017148 | 0.00098617 | 0.0014712 | 0.0019495 | 0.0012907 | 0.0016004 | 0.0015102 | 0.00031559 | 20.8978 | -0.001079 | -41.6741
| Zr | 0.015188 | 0.014954 | 0.015371 | 0.01568 | 0.014806 | 0.014807 | 0.015421 | 0.016051 | 0.015094 | 0.015735 | 0.015311 | 0.0004186 | 2.734 | -0.0018177 | -10.6122
| Nb | 0.0014695 | 0.0014995 | 0.001451 | 0.0017391 | 0.0015378 | 0.001487 | 0.0014337 | 0.0015955 | 0.0014757 | 0.0014131 | 0.0015102 | 9.585e-05 | 6.3469 | -0.00028231 | -15.7497
| Ba | 0.006754 | 0.012018 | 0.01534 | 0.0099048 | 0.013344 | 0.0099444 | 0.0087176 | 0.012334 | 0.0084395 | 0.014181 | 0.011098 | 0.0027743 | 24.9991 | -0.0018482 | -14.2764
| Th | 0.015828 | 0.015756 | 0.01587 | 0.015666 | 0.01595 | 0.015221 | 0.015969 | 0.015444 | 0.015617 | 0.015492 | 0.015681 | 0.00024156 | 1.5405 | 0.015562 | 13022.2567

## Table S7 T1-G
Non-calibrated data (wt%): T1-G
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) | Calibration factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 2.7939 | 3.3234 | 3.1356 | 2.9345 | 3.0989 | 2.9828 | 2.6605 | 2.81 | 2.9326 | 2.6439 | 2.9316 | 0.21492 | 7.3312 | -0.19845 | -6.3401 | 1.0677
| MgO | 2.8294 | 2.8507 | 2.8971 | 2.8519 | 2.9257 | 2.8679 | 2.9226 | 2.9116 | 2.8799 | 2.8743 | 2.8811 | 0.032622 | 1.1323 | -0.86895 | -23.1717 | 1.3016
| Al2O3 | 17.0094 | 16.9284 | 16.9365 | 16.9544 | 16.8902 | 16.9442 | 16.9424 | 16.9771 | 16.8764 | 17.0012 | 16.946 | 0.042792 | 0.25252 | -0.1543 | -0.90231 | 1.0091
| SiO2 | 57.8715 | 57.5528 | 57.6358 | 57.8117 | 57.6751 | 57.7564 | 57.9722 | 57.858 | 57.8661 | 57.9656 | 57.7965 | 0.13966 | 0.24164 | -0.80455 | -1.3729 | 1.0139
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.168 | -100 | nd
| K2O | 2.7416 | 2.7478 | 2.7238 | 2.7254 | 2.7161 | 2.6996 | 2.6832 | 2.6848 | 2.6882 | 2.6869 | 2.7097 | 0.024404 | 0.90059 | 0.7497 | 38.2492 | 0.72333
| CaO | 7.1582 | 7.1051 | 7.127 | 7.1513 | 7.1335 | 7.1557 | 7.1902 | 7.1614 | 7.1676 | 7.1923 | 7.1542 | 0.027037 | 0.37792 | 0.054108 | 0.76206 | 0.99244
| TiO2 | 0.80194 | 0.77717 | 0.79787 | 0.79811 | 0.8031 | 0.81173 | 0.8262 | 0.82108 | 0.82114 | 0.82484 | 0.80832 | 0.015574 | 1.9267 | 0.053304 | 7.06 | 0.93406
| V | 0.015375 | 0.014272 | 0.0192 | 0.019082 | 0.017823 | 0.015001 | 0.018948 | 0.020562 | 0.020211 | 0.022556 | 0.018303 | 0.0026794 | 14.6389 | -0.00069734 | -3.6701 | 1.0381
| Cr | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.00209 | -100 | nd
| MnO | 0.14973 | 0.14842 | 0.15094 | 0.15028 | 0.15083 | 0.1514 | 0.15288 | 0.15247 | 0.15244 | 0.15372 | 0.15131 | 0.0016054 | 1.061 | 0.02431 | 19.1412 | 0.83934
| Fe2O3t | 8.5599 | 8.4626 | 8.5072 | 8.5377 | 8.5237 | 8.5497 | 8.5991 | 8.5698 | 8.5551 | 8.6022 | 8.5467 | 0.041976 | 0.49114 | 1.3897 | 19.417 | 0.8374
| Co | 0.02936 | 0.031339 | 0.027913 | 0.028276 | 0.028185 | 0.027391 | 0.025566 | 0.025798 | 0.025874 | 0.02565 | 0.027535 | 0.0018897 | 6.863 | 0.025645 | 1356.8565 | 0.068641
| Ni | 0.00036384 | 0.0002275 | 0.00027773 | 0.00045305 | 0.00054196 | 0.00022287 | 0.00023764 | 0.00034903 | 0.00042747 | 0.00084389 | 0.0003945 | 0.00019056 | 48.3036 | -0.00066552 | -62.7839 | 2.687
| Cu | 0.0031597 | 0.0035187 | 0.0027222 | 0.0030188 | 0.0031248 | 0.0034777 | 0.0028342 | 0.0035281 | 0.0036293 | 0.0026469 | 0.003166 | 0.00036012 | 11.3745 | 0.001286 | 68.4033 | 0.59381
| Zn | 0.0090848 | 0.0085172 | 0.0083928 | 0.0086156 | 0.0084425 | 0.008964 | 0.0087519 | 0.0085443 | 0.0088547 | 0.0085662 | 0.0086734 | 0.00023111 | 2.6646 | 0.0012733 | 17.206 | 0.8532
| Ga | 0.0035575 | 0.0036326 | 0.0032706 | 0.0036377 | 0.003559 | 0.0035126 | 0.0033815 | 0.003225 | 0.0033477 | 0.0035105 | 0.0034635 | 0.0001475 | 4.2587 | 0.0015234 | 78.5261 | 0.56014
| Rb | 0.012473 | 0.01233 | 0.012058 | 0.012275 | 0.012193 | 0.011801 | 0.011976 | 0.012621 | 0.012336 | 0.012444 | 0.01225 | 0.00024869 | 2.03 | 0.0042803 | 53.7045 | 0.6506
| Sr | 0.037577 | 0.03644 | 0.036829 | 0.037411 | 0.037425 | 0.036733 | 0.036918 | 0.036984 | 0.037142 | 0.037662 | 0.037112 | 0.00039952 | 1.0765 | 0.0087114 | 30.6734 | 0.76527
| Y | 0.001924 | 0.0020252 | 0.0024001 | 0.00228 | 0.0021551 | 0.0024283 | 0.0025904 | 0.0021222 | 0.0021507 | 0.0020069 | 0.0022083 | 0.00021208 | 9.6038 | -0.00018176 | -7.6048 | 1.0823
| Zr | 0.031994 | 0.03187 | 0.032043 | 0.032633 | 0.03217 | 0.032828 | 0.032629 | 0.03184 | 0.031766 | 0.03278 | 0.032255 | 0.00041722 | 1.2935 | 0.017855 | 123.992 | 0.44644
| Nb | 0.0037073 | 0.003172 | 0.0039724 | 0.0039811 | 0.0032259 | 0.0036348 | 0.0042484 | 0.0039484 | 0.0037375 | 0.0037621 | 0.003739 | 0.0003356 | 8.9757 | 0.002852 | 321.5241 | 0.23723
| Ba | 0.122 | 0.12507 | 0.12444 | 0.11912 | 0.11727 | 0.12457 | 0.10078 | 0.09957 | 0.097055 | 0.089935 | 0.11198 | 0.013551 | 12.1014 | 0.073181 | 188.6061 | 0.34649
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0070401 | -100 | nd
| Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0127 | -100 | nd
| Th | 0.014677 | 0.01421 | 0.014208 | 0.014668 | 0.014736 | 0.015375 | 0.014459 | 0.014989 | 0.015053 | 0.013719 | 0.014609 | 0.00048138 | 3.295 | 0.011489 | 368.2431 | 0.21356
Calibrated data(wt%): T1-G
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 2.4419 | 2.9061 | 2.7367 | 2.5635 | 2.7022 | 2.6046 | 2.3175 | 2.449 | 2.5588 | 2.3061 | 2.5586 | 0.18938 | 7.4017 | -0.57142 | -18.2558
| MgO | 7.7127 | 7.7706 | 7.8841 | 7.7694 | 7.9554 | 7.8095 | 7.9423 | 7.9153 | 7.8375 | 7.8217 | 7.8418 | 0.080716 | 1.0293 | 4.0918 | 109.1118
| Al2O3 | 17.1425 | 17.0644 | 17.0571 | 17.0821 | 17.0044 | 17.0676 | 17.0479 | 17.0856 | 16.9925 | 17.1169 | 17.0661 | 0.045581 | 0.26709 | -0.034225 | -0.20015
| SiO2 | 55.8384 | 55.5455 | 55.5614 | 55.7595 | 55.5715 | 55.6903 | 55.8189 | 55.72 | 55.7632 | 55.8513 | 55.712 | 0.11662 | 0.20932 | -2.8891 | -4.9301
| P2O5 | 0.0018488 | 0.0018449 | 0.0018362 | 0.0018436 | 0.0018402 | 0.0018417 | 0.0018468 | 0.0018531 | 0.0018412 | 0.0018484 | 0.0018445 | 4.9406e-06 | 0.26786 | -0.16616 | -98.9021
| K2O | 1.9322 | 1.9373 | 1.9171 | 1.9197 | 1.9103 | 1.901 | 1.8853 | 1.8871 | 1.8913 | 1.8899 | 1.9071 | 0.019047 | 0.99875 | -0.05292 | -2.6999
| CaO | 6.9689 | 6.92 | 6.9298 | 6.9592 | 6.9308 | 6.9618 | 6.98 | 6.9548 | 6.9677 | 6.9901 | 6.9563 | 0.022842 | 0.32837 | -0.14382 | -2.0256
| TiO2 | 0.73329 | 0.71419 | 0.72938 | 0.72961 | 0.73124 | 0.73766 | 0.74916 | 0.74655 | 0.746 | 0.75113 | 0.73682 | 0.011545 | 1.5669 | -0.018192 | -2.4096
| V | 0.01313 | 0.010096 | 0.017098 | 0.015511 | 0.015206 | 0.01536 | 0.018173 | 0.019104 | 0.019651 | 0.018984 | 0.016231 | 0.0030117 | 18.5549 | -0.0027692 | -14.5743
| Cr | 0.0032363 | 0.0037603 | 0.0026527 | 0.002767 | 0.002337 | 0.0030095 | 0.0019543 | 0.0017971 | 0.0024683 | 0.0024032 | 0.0026386 | 0.00058892 | 22.3198 | 0.00054853 | 26.2448
| MnO | 0.11899 | 0.11732 | 0.11954 | 0.11907 | 0.11927 | 0.12082 | 0.12132 | 0.12106 | 0.12125 | 0.12173 | 0.12004 | 0.00141 | 1.1746 | -0.0069656 | -5.4846
| Fe2O3t | 6.9949 | 6.9172 | 6.9425 | 6.975 | 6.9516 | 6.9818 | 7.011 | 6.9916 | 6.9893 | 7.0256 | 6.9781 | 0.032709 | 0.46874 | -0.17896 | -2.5004
| Co | 0.0015691 | 0.0016761 | 0.0014888 | 0.0015102 | 0.0015023 | 0.0014626 | 0.0013609 | 0.0013747 | 0.0013809 | 0.0013679 | 0.0014693 | 0.00010258 | 6.9812 | -0.00042069 | -22.2581
| Ni | 0.00054624 | 0.00033614 | 0.00041486 | 0.00068393 | 0.00081845 | 0.00033116 | 0.00035441 | 0.00052495 | 0.00064588 | 0.0012841 | 0.00059401 | 0.0002924 | 49.2246 | -0.00046601 | -43.9625
| Cu | 0.0019731 | 0.0021979 | 0.0016975 | 0.0018847 | 0.0019469 | 0.0021703 | 0.0017652 | 0.0021988 | 0.0022655 | 0.0016515 | 0.0019751 | 0.00022533 | 11.4081 | 9.5102e-05 | 5.0585
| Zn | 0.0068614 | 0.0064344 | 0.0063295 | 0.0065054 | 0.0063617 | 0.0067655 | 0.0065922 | 0.0064407 | 0.0066854 | 0.0064645 | 0.0065441 | 0.00017698 | 2.7044 | -0.00085606 | -11.5682
| Ga | 0.0018408 | 0.0018802 | 0.0016899 | 0.001882 | 0.0018374 | 0.0018164 | 0.001745 | 0.0016656 | 0.0017317 | 0.0018151 | 0.0017904 | 7.721e-05 | 4.3124 | -0.00014961 | -7.7119
| Rb | 0.0074921 | 0.0074077 | 0.0072322 | 0.0073718 | 0.0073072 | 0.0070837 | 0.0071749 | 0.0075672 | 0.007409 | 0.0074702 | 0.0073516 | 0.00015141 | 2.0595 | -0.00061854 | -7.7607
| Sr | 0.026777 | 0.025972 | 0.026206 | 0.026653 | 0.026607 | 0.026158 | 0.02624 | 0.026308 | 0.026464 | 0.026822 | 0.026421 | 0.00028684 | 1.0857 | -0.0019799 | -6.9714
| Y | 0.0023637 | 0.0024885 | 0.0029459 | 0.0028016 | 0.0026421 | 0.0029833 | 0.0031766 | 0.0026031 | 0.0026429 | 0.0024644 | 0.0027112 | 0.00025956 | 9.5735 | 0.00032117 | 13.4378
| Zr | 0.018018 | 0.017951 | 0.018019 | 0.018374 | 0.018076 | 0.018476 | 0.018329 | 0.017899 | 0.017889 | 0.01845 | 0.018148 | 0.00023308 | 1.2843 | 0.0037476 | 26.0247
| Nb | 0.0012861 | 0.0011009 | 0.0013758 | 0.0013806 | 0.0011168 | 0.0012601 | 0.0014696 | 0.0013671 | 0.0012964 | 0.0013043 | 0.0012958 | 0.00011563 | 8.9238 | 0.00040875 | 46.0815
| Ba | 0.016637 | 0.011863 | 0.021121 | 0.021052 | 0.023004 | 0.022769 | 0.023921 | 0.025292 | 0.026942 | 0.021209 | 0.021381 | 0.004353 | 20.3591 | -0.01742 | -44.8954
| Th | 0.0077747 | 0.0074281 | 0.0073661 | 0.0071209 | 0.0068968 | 0.007134 | 0.0072871 | 0.0071302 | 0.0073107 | 0.0071224 | 0.0072571 | 0.00023818 | 3.282 | 0.0055626 | 328.2608

## Table S8 GOR-132
Non-calibrated data (wt%): GOR-132
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) | Calibration factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 0 | 0.15962 | 0 | 0 | 0.10696 | 0 | 0 | 0 | 0.012269 | 0.14935 | 0.04282 | 0.067527 | 157.7013 | -0.78415 | -94.8221 | 19.3127
| MgO | 21.9871 | 21.9695 | 21.9823 | 21.9627 | 22.0615 | 21.9202 | 22.0323 | 21.9879 | 22.0083 | 22.0303 | 21.9942 | 0.040553 | 0.18438 | -0.32391 | -1.4513 | 1.0147
| Al2O3 | 9.6679 | 9.6926 | 9.7258 | 9.6563 | 9.6828 | 9.6446 | 9.7274 | 9.6212 | 9.6776 | 9.6389 | 9.6735 | 0.035294 | 0.36486 | -1.2863 | -11.7364 | 1.133
| SiO2 | 44.3821 | 44.2539 | 44.3908 | 44.4255 | 44.2732 | 44.4724 | 44.3399 | 44.3924 | 44.437 | 44.3792 | 44.3746 | 0.068915 | 0.1553 | -0.95907 | -2.1156 | 1.0216
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.035868 | -100 | nd
| K2O | 0.13443 | 0.13464 | 0.13491 | 0.13887 | 0.14022 | 0.13534 | 0.13886 | 0.13903 | 0.13859 | 0.13632 | 0.13712 | 0.0022014 | 1.6055 | 0.10643 | 346.8283 | 0.2238
| CaO | 8.6085 | 8.5842 | 8.6112 | 8.6148 | 8.5719 | 8.636 | 8.5802 | 8.6202 | 8.5847 | 8.5657 | 8.5977 | 0.023362 | 0.27173 | 0.17862 | 2.1216 | 0.97922
| TiO2 | 0.33431 | 0.32752 | 0.32996 | 0.33583 | 0.33072 | 0.32952 | 0.33222 | 0.33072 | 0.3309 | 0.33026 | 0.3312 | 0.0023909 | 0.7219 | 0.026314 | 8.6309 | 0.92055
| V | 0.029322 | 0.029798 | 0.031275 | 0.028482 | 0.028216 | 0.031988 | 0.030882 | 0.030247 | 0.031268 | 0.030573 | 0.030205 | 0.0012438 | 4.1177 | 0.0088832 | 41.6627 | 0.7059
| Cr | 0.28404 | 0.2834 | 0.28359 | 0.28528 | 0.28297 | 0.28501 | 0.28261 | 0.2844 | 0.28256 | 0.28394 | 0.28378 | 0.00093967 | 0.33113 | 0.031905 | 12.6669 | 0.88757
| MnO | 0.17619 | 0.17434 | 0.17535 | 0.17568 | 0.17441 | 0.17728 | 0.17607 | 0.17602 | 0.17483 | 0.17538 | 0.17555 | 0.00089815 | 0.51161 | 0.022118 | 14.4149 | 0.87401
| Fe2O3t | 14.1465 | 14.1176 | 14.1627 | 14.1775 | 14.0718 | 14.2086 | 14.1057 | 14.1551 | 14.0805 | 14.0339 | 14.126 | 0.053801 | 0.38087 | 2.9427 | 26.3134 | 0.79168
| Co | 0.052772 | 0.052448 | 0.053271 | 0.052681 | 0.052297 | 0.052794 | 0.052129 | 0.052916 | 0.05161 | 0.052406 | 0.052532 | 0.00046333 | 0.88199 | 0.043296 | 468.7712 | 0.17582
| Ni | 0.14145 | 0.141 | 0.14067 | 0.14176 | 0.14093 | 0.14168 | 0.14059 | 0.14008 | 0.13953 | 0.14115 | 0.14088 | 0.00070117 | 0.4977 | 0.022617 | 19.1239 | 0.83946
| Cu | 0.027994 | 0.028226 | 0.027806 | 0.028011 | 0.028037 | 0.028169 | 0.027824 | 0.028089 | 0.027905 | 0.027515 | 0.027958 | 0.00020665 | 0.73915 | 0.0075325 | 36.8787 | 0.73057
| Zn | 0.006847 | 0.0067996 | 0.0065647 | 0.0066565 | 0.0066581 | 0.0061805 | 0.0069327 | 0.0064841 | 0.0067719 | 0.006184 | 0.0066079 | 0.00026049 | 3.9421 | -0.001044 | -13.6438 | 1.158
| Ga | 0.0021065 | 0.001784 | 0.0016642 | 0.0020495 | 0.0018615 | 0.0016898 | 0.0017122 | 0.0015648 | 0.0015603 | 0.001744 | 0.0017737 | 0.00018484 | 10.4214 | 0.00073747 | 71.1706 | 0.58421
| Rb | 0.0087771 | 0.0083116 | 0.0082785 | 0.0089405 | 0.0083956 | 0.0082324 | 0.0085308 | 0.0085635 | 0.0085853 | 0.0080281 | 0.0084643 | 0.00027081 | 3.1995 | 0.0082551 | 3945.4283 | 0.024719
| Sr | 0.0031386 | 0.0027796 | 0.003659 | 0.0034217 | 0.0034424 | 0.0027975 | 0.0028877 | 0.0035008 | 0.0030487 | 0.003515 | 0.0032191 | 0.00032788 | 10.1856 | 0.0016947 | 111.1702 | 0.47355
| Y | 0 | 0.00032193 | 0 | 0 | 0 | 0 | 0.0001296 | 0 | 0.00018578 | 0.00036196 | 9.9927e-05 | 0.0001438 | 143.907 | -0.0011854 | -92.2253 | 12.8622
| Zr | 0.019373 | 0.019056 | 0.019097 | 0.019789 | 0.019133 | 0.019967 | 0.019669 | 0.019705 | 0.019264 | 0.019632 | 0.019468 | 0.00032431 | 1.6658 | 0.018482 | 1873.7179 | 0.050666
| Nb | 0.0039924 | 0.0040809 | 0.0038337 | 0.004324 | 0.0037363 | 0.0037996 | 0.0042222 | 0.0043332 | 0.0038957 | 0.0040995 | 0.0040318 | 0.00021603 | 5.3581 | 0.0040245 | 55332.0671 | 0.001804
| Ba | 0.12809 | 0.12575 | 0.12624 | 0.13357 | 0.12724 | 0.13462 | 0.11142 | 0.11389 | 0.11598 | 0.10565 | 0.12224 | 0.0098353 | 8.0456 | 0.12216 | 150443.5544 | 0.00066426
| La | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -8.3892e-06 | -100 | nd
| Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -3.9156e-05 | -100 | nd
| Th | 0.021039 | 0.022894 | 0.021464 | 0.02235 | 0.022138 | 0.02186 | 0.021437 | 0.02201 | 0.02283 | 0.020617 | 0.021864 | 0.00073829 | 3.3768 | 0.021863 | 2438122.842 | 4.1013e-05
Calibrated data (wt%): GOR-132
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 0.09102 | 0.0074797 | 0 | 0 | 0 | 0.065144 | 0 | 0 | 0.097404 | 0 | 0.026105 | 0.041173 | 157.7237 | -0.80086 | -96.8433
| MgO | 42.5573 | 42.5295 | 42.5027 | 42.5623 | 42.4123 | 42.6097 | 42.4726 | 42.4937 | 42.4858 | 42.5032 | 42.5129 | 0.054872 | 0.12907 | 20.1948 | 90.4861
| Al2O3 | 7.9499 | 7.9842 | 7.9408 | 8.0227 | 7.9653 | 7.9858 | 7.9718 | 8.0264 | 8.0018 | 7.9798 | 7.9829 | 0.028187 | 0.35309 | -2.9769 | -27.1622
| SiO2 | 34.4358 | 34.4889 | 34.4726 | 34.3998 | 34.5604 | 34.3519 | 34.5085 | 34.4636 | 34.3726 | 34.4646 | 34.4519 | 0.063538 | 0.18442 | -10.8818 | -24.0038
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.035868 | -100
| K2O | 0.072212 | 0.073423 | 0.073696 | 0.073545 | 0.071821 | 0.074286 | 0.073653 | 0.071522 | 0.071427 | 0.071277 | 0.072686 | 0.00114 | 1.5684 | 0.041999 | 136.8598
| CaO | 6.3063 | 6.3208 | 6.3505 | 6.3157 | 6.3686 | 6.3105 | 6.3491 | 6.3441 | 6.3287 | 6.3425 | 6.3337 | 0.020357 | 0.32141 | -2.0854 | -24.7703
| TiO2 | 0.21853 | 0.21847 | 0.2191 | 0.22012 | 0.21936 | 0.21975 | 0.22269 | 0.21961 | 0.21754 | 0.2208 | 0.2196 | 0.001426 | 0.64936 | -0.085284 | -27.9728
| V | 0.021787 | 0.021242 | 0.021173 | 0.021583 | 0.021752 | 0.019917 | 0.020535 | 0.021971 | 0.02049 | 0.020493 | 0.021094 | 0.00069708 | 3.3046 | -0.0002276 | -1.0675
| Cr | 0.1929 | 0.19184 | 0.19229 | 0.1921 | 0.19418 | 0.19214 | 0.19338 | 0.19304 | 0.19407 | 0.19337 | 0.19293 | 0.00082853 | 0.42945 | -0.058945 | -23.4025
| MnO | 0.096067 | 0.09559 | 0.096591 | 0.096313 | 0.097024 | 0.0956 | 0.09651 | 0.095982 | 0.095266 | 0.096449 | 0.096139 | 0.00054113 | 0.56286 | -0.057298 | -37.343
| Fe2O3 | 7.9833 | 8.0053 | 8.0509 | 8.0183 | 8.085 | 7.9956 | 8.0622 | 8.0536 | 8.0359 | 8.0438 | 8.0334 | 0.032094 | 0.39951 | -3.1499 | -28.1664
| Co | 0.0018662 | 0.0018352 | 0.0018834 | 0.0018539 | 0.0018797 | 0.0018593 | 0.0018749 | 0.001897 | 0.0018681 | 0.0018784 | 0.0018696 | 1.7318e-05 | 0.92628 | -0.0073665 | -79.7577
| Ni | 0.13969 | 0.13799 | 0.13856 | 0.13899 | 0.14024 | 0.1393 | 0.14024 | 0.13914 | 0.13962 | 0.1399 | 0.13937 | 0.00072328 | 0.51897 | 0.021102 | 17.8426
| Cu | 0.011048 | 0.011198 | 0.011274 | 0.011161 | 0.011314 | 0.011244 | 0.011243 | 0.011159 | 0.01134 | 0.011233 | 0.011221 | 8.4904e-05 | 0.75663 | -0.0092038 | -45.0611
| Zn | 0.0029985 | 0.0032819 | 0.003143 | 0.0033584 | 0.0029978 | 0.003225 | 0.0032269 | 0.0031817 | 0.0032992 | 0.0033182 | 0.0032031 | 0.00012559 | 3.9208 | -0.0044489 | -58.1404
| Ga | 0.000578 | 0.00051677 | 0.00051835 | 0.00056682 | 0.00056022 | 0.0006161 | 0.00067893 | 0.00055127 | 0.0005915 | 0.00069754 | 0.00058755 | 6.1157e-05 | 10.4088 | -0.00044865 | -43.2975
| Rb | 0.0030442 | 0.0032531 | 0.0032456 | 0.0032314 | 0.0031223 | 0.0031794 | 0.0033886 | 0.0031372 | 0.0031534 | 0.0033258 | 0.0032081 | 0.00010191 | 3.1767 | 0.0029989 | 1433.2762
| Sr | 0.0015787 | 0.0013682 | 0.0015713 | 0.0012955 | 0.0012567 | 0.0015439 | 0.0015361 | 0.0016424 | 0.0012489 | 0.0014085 | 0.001445 | 0.00014695 | 10.1693 | -7.9375e-05 | -5.2069
| Y | 0.00027211 | 0.00013491 | 0 | 9.1208e-05 | 0 | 0 | 0 | 0 | 0.00024062 | 0 | 7.3884e-05 | 0.00010743 | 145.403 | -0.0012114 | -94.2515
| Zr | 0.0069479 | 0.0068126 | 0.00697 | 0.006953 | 0.0070677 | 0.0067621 | 0.0069997 | 0.0067537 | 0.0067475 | 0.0068506 | 0.0068865 | 0.00011547 | 1.6767 | 0.0059001 | 598.1551
| Nb | 0.00089461 | 0.00084964 | 0.00094499 | 0.00092038 | 0.0008297 | 0.00081458 | 0.0009431 | 0.00083641 | 0.00089106 | 0.00087079 | 0.00087953 | 4.7007e-05 | 5.3446 | 0.00087225 | 11992.5119
| Ba | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -8.1202e-05 | -100
| Th | 0.011295 | 0.012495 | 0.012052 | 0.011731 | 0.011979 | 0.012109 | 0.012235 | 0.011753 | 0.012545 | 0.011517 | 0.011971 | 0.00040365 | 3.3719 | 0.01197 | 1334895.0414

## Table S9 OOL-31A
Non-calibrated data (wt%): OOL-31A
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation) | Calibration factor
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 4.8306 | 5.0258 | 5.001 | 5.1947 | 5.0635 | 4.681 | 4.9894 | 5.0652 | 4.7021 | 4.9207 | 4.9474 | 0.16492 | 3.3335 | 1.2094 | 32.3547 | 0.75555
| MgO | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | 0 | nd | nd
| Al2O3 | 12.3482 | 12.2861 | 12.2979 | 12.2723 | 12.3169 | 12.3393 | 12.3235 | 12.2836 | 12.3439 | 12.2839 | 12.3096 | 0.028302 | 0.22992 | -0.43945 | -3.447 | 1.0357
| SiO2 | 74.9491 | 74.8616 | 74.8375 | 74.7148 | 74.7797 | 75.0997 | 74.8317 | 74.7799 | 75.0357 | 74.9192 | 74.8809 | 0.12049 | 0.16091 | -1.1846 | -1.5574 | 1.0158
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.48843 | -100 | nd
| K2O | 6.6513 | 6.6316 | 6.6298 | 6.6051 | 6.6162 | 6.6588 | 6.6228 | 6.6191 | 6.6655 | 6.6354 | 6.6336 | 0.019528 | 0.29438 | 1.1811 | 21.6616 | 0.82195
| CaO | 0.40995 | 0.40974 | 0.41748 | 0.4061 | 0.40923 | 0.41135 | 0.41117 | 0.41133 | 0.41172 | 0.4116 | 0.41097 | 0.0028477 | 0.69293 | 0.0022814 | 0.55823 | 0.99445
| TiO2 | 0.099657 | 0.099699 | 0.10015 | 0.098224 | 0.098364 | 0.099199 | 0.097426 | 0.09939 | 0.09898 | 0.096757 | 0.098784 | 0.0010792 | 1.0925 | -0.010863 | -9.9073 | 1.11
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0011962 | -100 | nd
| Cr | 0 | 4.5373e-05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.5373e-06 | 1.4348e-05 | 316.2278 | -0.00039418 | -98.862 | 87.8763
| MnO | 0.12919 | 0.12723 | 0.12886 | 0.12725 | 0.12789 | 0.12918 | 0.12817 | 0.12797 | 0.13012 | 0.12891 | 0.12848 | 0.00092907 | 0.72314 | 0.018831 | 17.1737 | 0.85343
| Fe2O3t | 0.8426 | 0.83785 | 0.84133 | 0.8374 | 0.8366 | 0.84615 | 0.83839 | 0.8398 | 0.84508 | 0.84162 | 0.84068 | 0.0032622 | 0.38804 | 0.02331 | 2.8519 | 0.97227
| Co | 0.0026106 | 0.0023435 | 0.0022843 | 0.0027314 | 0.0023975 | 0.0026346 | 0.0026158 | 0.0022966 | 0.0024733 | 0.002545 | 0.0024933 | 0.00015721 | 6.3056 | nd | nd | nd
| Ni | 0.00066397 | 0.00064592 | 0.00048504 | 0.0006147 | 0.00059329 | 0.00071873 | 0.00047349 | 0.00045304 | 0.00068581 | 0.00046793 | 0.00058019 | 0.00010125 | 17.4506 | nd | nd | nd
| Cu | 0.0014649 | 0.0012581 | 0.0011834 | 0.0014526 | 0.001217 | 0.0013055 | 0.0016419 | 0.0011591 | 0.001322 | 0.001177 | 0.0013182 | 0.00015698 | 11.9088 | nd | nd | nd
| Zn | 0.00729 | 0.0069531 | 0.0068286 | 0.0070867 | 0.0071549 | 0.00733 | 0.0070988 | 0.0072581 | 0.0071917 | 0.0071919 | 0.0071384 | 0.00015441 | 2.1631 | 0.001357 | 23.4711 | 0.80991
| Ga | 0.0033892 | 0.0033105 | 0.0034084 | 0.0034028 | 0.0033233 | 0.0034771 | 0.0035682 | 0.0033925 | 0.0034177 | 0.0035431 | 0.0034233 | 8.4149e-05 | 2.4581 | 0.0014297 | 71.7139 | 0.58236
| Rb | 0.023701 | 0.023939 | 0.024034 | 0.023903 | 0.024023 | 0.023944 | 0.02366 | 0.024053 | 0.024268 | 0.023917 | 0.023944 | 0.0001743 | 0.72794 | 0.0034103 | 16.6081 | 0.85757
| Sr | 0.00072306 | 0.00098908 | 0.00082412 | 0.0010663 | 0.0010322 | 0.0010011 | 0.00099199 | 0.00091223 | 0.00099466 | 0.00087858 | 0.00094133 | 0.00010609 | 11.2699 | 0.00094133 | nd | 0
| Y | 0.0017434 | 0.001761 | 0.0016785 | 0.0016359 | 0.0021627 | 0.0018775 | 0.001798 | 0.0019509 | 0.0016014 | 0.0019144 | 0.0018124 | 0.00016922 | 9.3367 | -0.00048026 | nd | 1.265
| Zr | 0.021658 | 0.021632 | 0.022421 | 0.021781 | 0.02165 | 0.022004 | 0.021637 | 0.021808 | 0.021933 | 0.02174 | 0.021826 | 0.00024475 | 1.1213 | 0.0087684 | 67.1497 | 0.59827
| Nb | 0.0052423 | 0.0051951 | 0.0055026 | 0.0056865 | 0.0053101 | 0.0054522 | 0.0049779 | 0.0053676 | 0.005623 | 0.0053194 | 0.0053677 | 0.00020954 | 3.9038 | 0.0022776 | 73.7072 | 0.57568
| Ba | 0.014281 | 0 | 0.021309 | 0.0063576 | 0.015611 | 0.015313 | 0.012922 | 0.027862 | 0.016904 | 0.020145 | 0.015071 | 0.007748 | 51.4114 | 0.014273 | 1789.8794 | 0.052913
| La | 0.012806 | 0.015564 | 0.012995 | 0.012781 | 0.01526 | 0.013091 | 0.015925 | 0.018057 | 0.014246 | 0.017313 | 0.014804 | 0.0019297 | 13.0355 | 0.011116 | 301.387 | 0.24914
| Ce | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0066785 | -100 | nd
| Th | 0.0077747 | 0.0074281 | 0.0073661 | 0.0071209 | 0.0068968 | 0.007134 | 0.0072871 | 0.0071302 | 0.0073107 | 0.0071224 | 0.0072571 | 0.00023818 | 3.282 | 0.0055626 | 328.2608 | 0.2335
Calibrated data (wt%): OOL-31A
|  | Analysis number | [11 empty] | Coeff. Variation | Absolute deviation (wt%) | % relative error (deviation)
| Element | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean (n=10) | ± 1 σ | %
| Na2O | 4.7427 | 4.5297 | 4.8823 | 4.8077 | 4.5112 | 4.8798 | 5.006 | 4.8207 | 4.8436 | 4.6549 | 4.7679 | 0.15934 | 3.342 | 1.0299 | 27.5518
| MgO | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | 0 | nd
| Al2O3 | 12.9412 | 12.9987 | 12.9406 | 12.9798 | 12.9999 | 12.9743 | 12.9269 | 12.9575 | 12.9438 | 13.0077 | 12.9671 | 0.028968 | 0.22339 | 0.21803 | 1.7102
| SiO2 | 76.4225 | 76.5054 | 76.2808 | 76.3132 | 76.6073 | 76.2697 | 76.1965 | 76.3481 | 76.3592 | 76.4451 | 76.3748 | 0.12235 | 0.16019 | 0.30925 | 0.40656
| P2O5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.48843 | -100
| K2O | 5.0584 | 5.0785 | 5.0465 | 5.0472 | 5.0756 | 5.0429 | 5.034 | 5.0548 | 5.0544 | 5.0695 | 5.0562 | 0.014567 | 0.28811 | -0.39631 | -7.2684
| CaO | 0.42704 | 0.42689 | 0.4268 | 0.42648 | 0.42668 | 0.42453 | 0.42126 | 0.43319 | 0.42505 | 0.42519 | 0.42631 | 0.0029839 | 0.69992 | 0.017624 | 4.3124
| TiO2 | 0.093908 | 0.095854 | 0.095556 | 0.094628 | 0.09576 | 0.095084 | 0.094725 | 0.096507 | 0.097301 | 0.096173 | 0.09555 | 0.00099721 | 1.0437 | -0.014098 | -12.8575
| V | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | nd | -0.0011962 | -100
| Cr | 0.00029639 | 0.00028563 | 0.00016371 | 0.00032172 | 0.0002436 | 0.00048498 | 0.00027974 | 0.00034634 | 0.00067372 | 0.00043324 | 0.00035291 | 0.00014478 | 41.0249 | -4.5811e-05 | -11.4896
| MnO | 0.1127 | 0.11358 | 0.11199 | 0.11196 | 0.11262 | 0.11176 | 0.11118 | 0.11269 | 0.11115 | 0.11284 | 0.11225 | 0.00077477 | 0.69022 | 0.0026009 | 2.372
| Fe2O3 | 0.74875 | 0.75167 | 0.74686 | 0.74605 | 0.7531 | 0.74452 | 0.74554 | 0.74867 | 0.74603 | 0.74986 | 0.74811 | 0.0028077 | 0.3753 | -0.069267 | -8.4744
| Co | 0.00014894 | 0.00014454 | 0.00013435 | 0.00015315 | 0.00015419 | 0.00014024 | 0.00016024 | 0.00013363 | 0.00013738 | 0.00015286 | 0.00014595 | 9.2772e-06 | 6.3563 | nd | nd
| Ni | 0.00080234 | 0.0011752 | 0.00077661 | 0.00081204 | 0.0012323 | 0.0010174 | 0.0010545 | 0.00083188 | 0.0011083 | 0.0011385 | 0.00099491 | 0.00017357 | 17.4463 | nd | nd
| Cu | 0.0008219 | 0.00092283 | 0.00080923 | 0.0011468 | 0.00091193 | 0.00085008 | 0.0010152 | 0.00082667 | 0.00087954 | 0.0010232 | 0.00092074 | 0.00010971 | 11.9157 | nd | nd
| Zn | 0.0060783 | 0.0060762 | 0.0061328 | 0.006001 | 0.0061971 | 0.006049 | 0.0059949 | 0.0057733 | 0.0058834 | 0.0061632 | 0.0060349 | 0.00012922 | 2.1412 | 0.0002535 | 4.3847
| Ga | 0.0020529 | 0.0019797 | 0.0019652 | 0.002068 | 0.0020155 | 0.0019263 | 0.0019735 | 0.0019756 | 0.0019206 | 0.0019645 | 0.0019842 | 4.8395e-05 | 2.439 | -9.4089e-06 | -0.47196
| Rb | 0.016143 | 0.016375 | 0.016229 | 0.015974 | 0.016168 | 0.016221 | 0.016152 | 0.016228 | 0.016181 | 0.016004 | 0.016167 | 0.00011491 | 0.71075 | -0.0043665 | -21.2648
| Sr | 0.0007038 | 0.00079657 | 0.00073051 | 0.00079493 | 0.00080233 | 0.00082727 | 0.00085518 | 0.00066042 | 0.00079351 | 0.00057952 | 0.0007544 | 8.5155e-05 | 11.2877 | 0.0007544 | nd
| Y | 0.0026436 | 0.0022092 | 0.0026932 | 0.0024837 | 0.0025938 | 0.0029893 | 0.0022605 | 0.0023176 | 0.0024354 | 0.0024082 | 0.0025034 | 0.00023431 | 9.3595 | 0.00021081 | 9.1953
| Zr | 0.013769 | 0.013889 | 0.013808 | 0.01371 | 0.013944 | 0.01372 | 0.013812 | 0.014207 | 0.013723 | 0.013725 | 0.013831 | 0.00015332 | 1.1085 | 0.00077269 | 5.9173
| Nb | 0.0020742 | 0.0021919 | 0.0020922 | 0.0019419 | 0.002127 | 0.0020715 | 0.0022199 | 0.0021464 | 0.002029 | 0.0020452 | 0.0020939 | 8.1591e-05 | 3.8965 | -0.00099614 | -32.2369
| Ba | 0 | 0 | 0.0018675 | 0 | 0.0014005 | 0 | 0 | 0 | 0 | 0 | 0.0003268 | 0.0006977 | 213.4922 | -0.00047063 | -59.0181
| Th | 0.0069754 | 0.0071585 | 0.0069816 | 0.0071382 | 0.0069902 | 0.0067593 | 0.0069816 | 0.0072171 | 0.0072836 | 0.0076163 | 0.0071102 | 0.00023323 | 3.2802 | 0.0054156 | 319.5899

## Table S10 Calibration Factors
Type-calibration factors used in the second stage quantification procedure
| Element | Calibration factor (published / measured)
| Na2O | 0.94 | (based on AGV-2)
| MgO | 2.94 | (based on AGV-2)
| Al2O3 | 1.03 | (based on AGV-2)
| SiO2 | 0.99 | (based on AGV-2)
| P2O5 | - | (based on AGV-2)
| K2O | 0.74 | (based on AGV-2)
| CaO | 1.04 | (based on AGV-2)
| TiO2 | 0.98 | (based on AGV-2)
| V | 1.38 | (based on AGV-2)
| Cr | 0.96 | (based on BHVO-2)
| MnO | 0.88 | (based on AGV-2)
| Fe2O3 | 0.9 | (based on AGV-2)
| Co | 0.06 | (based on AGV-2)
| Ni | 1.74 | (based on AGV-2)
| Cu | 0.71 | (based on AGV-2)
| Zn | 0.86 | (based on AGV-2)
| Ga | 0.59 | (based on AGV-2)
| Rb | 0.69 | (based on AGV-2)
| Sr | 0.82 | (based on AGV-2)
| Y | 1.42 | (based on AGV-2)
| Zr | 0.65 | (based on AGV-2)
| Nb | 0.4 | (based on AGV-2)
| Ba | 0.58 | (based on AGV-2)
| Th | 0.06 | (based on AGV-2)