APPENDICES

Appendix A – Hereford UK Raw Data and Calculations

| Details |  | Hereford |  | “Leptorapide” Results |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ID No. | Serogroup | IgM | MAT | 13/06/2006 |  | 14/06/2006 |  |  |
|  |  |  |  | Test 1 | Test 2 | Test 3 | Test 4 | Overall |
| 1 | - | 80 | -ve | -ve | -ve | -ve | -ve | -ve |
| 2 | Ictero | 640 | 320 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 3 | Hardjo | 160 | 320 | 3/5 | 5/5 | 4/5 | 3/5 | +ve |
| 4 | ND | 1280 | 1280 | 3/5 | 5/5 | 5/5 | 4/5 | +ve |
| 5 | ND | 2560 | 640 | 1/5 | 5/5 | 5/5 | 5/5 | +ve |
| 6 | ND | 1280 | 640 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 7 | Ictero | 1280 | 1280 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 8 | ND | 2560 | 640 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 9 | ND | 1280 | 320 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 10 | ND | 320 | 40 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 11 | Ictero | 160 | 160 | x | 4/5 | 3/5 | 4/5 | +ve |
| 12 | ND | 320 | 1280 | 2/5 | x | 4/5 | 5/5 | +ve |
| 13 | ND | 320 | 640 | 3/5 | x | 4/5 | 5/5 | +ve |
| 14 | ND | 640 | 320 | 4/5 | 2/5 | 4/5 | 5/5 | +ve |
| 15 | Saxkoebing | 320 | 640 | 3/5 | x | 5/5 | 5/5 | +ve |
| 16 | Hardjo | 640 | 5120 | 4/5 | x | 5/5 | 5/5 | +ve |
| 17 | ND | 2560 | 5120 | 4/5 | x | 4/5 | 4/5 | +ve |
| 18 | ND | 1280 | 2560 | 3/5 | x | 5/5 | 5/5 | +ve |
| 19 | ND | 2560 | 5120 | 4/5 | 3/5 | 5/5 | 5/5 | +ve |
| 20 | Ictero | 640 | 640 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 21 | ND | 2560 | 160 | 3/5 | 5/5 | 5/5 | 5/5 | +ve |
| 22 | - | 1280 | 160 | 3/5 | 2/5 | 2/5 | 5/5 | +ve |
| 23 | - | 160 | -ve | 3/5 | 3/5 | 5/5 | 5/5 | +ve |
| 24 | Saxkoebing | 80 | -ve | 2/5 | 4/5 | 1/5 | 2/5 | -ve |
| 25 | - | 160 | 1280 | 1/5 | 5/5 | 4/5 | 4/5 | +ve |
| 26 | ND | 160 | 320 | 2/5 | 5/5 | 4/5 | 4/5 | +ve |
| 27 | - | 80 | -ve | 2/5 | 3/5 | 2/5 | 2/5 | -ve |
| 28 | Ictero | -ve | 1280 | 1/5 | 2/5 | 1/5 | 2/5 | -ve |
| 29 | ND | 320 | 160 | 2/5 | 5/5 | 5/5 | 5/5 | +ve |
| 30 | Javanica | 160 | 1280 | 1/5 | 1/5 | 3/5 | 5/5 | +ve |
| 31 |  | -ve | -ve | 2/5 | 1/5 | 1/5 | 1/5 | -ve |
| 32 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 33 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 34 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 35 |  | -ve | -ve | 1/5 | 5/5 | 3/5 | 2/5 | +ve |
| 36 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 2/5 | -ve |
| 37 |  | -ve | -ve | 2/5 | 2/5 | 1/5 | 1/5 | -ve |
| 38 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 39 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 40 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 41 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 42 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 43 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 2/5 | -ve |
| 44 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 45 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 46 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 2/5 | -ve |
| 47 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 48 |  | -ve | -ve | 2/5 | 3/5 | 2/5 | 2/5 | -ve |
| 49 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 50 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 51 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 52 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 53 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 54 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |
| 55 |  | -ve | -ve | 1/5 | 1/5 | 1/5 | 1/5 | -ve |

Key:

X = Not smeared properly

1/5 = Negative (-ve)

2/5 = Inconclusive

3,4,5/5 = Positive (+ve)

Leptorapide” Evaluation

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| "Leptorapide" Positive | 2 | 25 | 27 |
| "Leptorapide" Negative | 27 | 1 | 28 |
| Totals | 29 | 26 | 55 |

|  | Estimated Value | 95% Confidence Interval |  |  |
| --- | --- | --- | --- | --- |
|  |  | Lower Limit | Upper Limit |  |
| Prevalence | 0.472727 | 0.338572 | 0.610681 |  |
| Sensitivity | 0.961538 | 0.784177 | 0.997989 |  |
| Specificity | 0.931034 | 0.757889 | 0.987964 |  |
| For any particular test result, the probability that it will be: |  |  |  |  |
| Positive | 0.490909 | 0.35533 |  | 0.627754 |
| Negative | 0.509091 | 0.372246 |  | 0.64467 |
| For any particular positive test result, the probability that it is: |  |  |  |  |
| True Positive | 0.925926 | 0.74247 |  | 0.987065 |
| False Positive | 0.074074 | 0.012935 |  | 0.25753 |
| For any particular negative test result, the probability that it is: |  |  |  |  |
| True Negative | 0.964286 | 0.79761 |  | 0.998133 |
| False Negative | 0.035714 | 0.001867 |  | 0.20239 |
| likelihood Ratios: |  |  |  |  |
| [C] = conventional |  |  |  |  |
| [W] = weighted by prevalence |  |  |  |  |
| Positive [C] | 13.942308 | 3.652685 |  | 53.21783 |
| Negative [C] | 0.041311 | 0.006015 |  | 0.28373 |
| Positive [W] | 12.5 | 3.280133 |  | 47.635257 |
| Negative [W] | 0.037037 | 0.005391 |  | 0.254462 |

Appendix B – K.I.T Netherlands Raw Data and Calculations

|  | K.I.T |  | “Leptorapide” |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| ID Number | IgM | MAT | Test 1 | Test 2 | Test 3 | Overall |
| 1 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 2 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 3 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 4 | 80 | 40 | 1/5 | 1/5 | 1/5 | -ve |
| 5 | 320 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 6 | 160 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 7 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 8 | 160 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 9 | 160 | 0 | 2/5 | 2/5 | 2/5 | -ve |
| 10 | 80 | 20 | 2/5 | 1/5 | 2/5 | -ve |
| 11 | 160 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 12 | 160 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 13 | 160 | 0 | 2/5 | 2/5 | 2/5 | -ve |
| 14 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 15 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 16 | 0 | 80 | 1/5 | 1/5 | 1/5 | -ve |
| 17 | 40 | 0 | 1/5 | 2/5 | 3/5 | -ve |
| 18 | 40 | 20 | 1/5 | 1/5 | 1/5 | -ve |
| 19 | 320 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 20 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 21 | 80 | 0 | 3/5 | 1/5 | 3/5 | +ve |
| 22 | 160 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 23 | 40 | 0 | 2/5 | 1/5 | 1/5 | -ve |
| 24 | 80 | 40 | 1/5 | 1/5 | 1/5 | -ve |
| 25 | 160 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 26 | 40 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 27 | 40 | 0 | 2/5 | 2/5 | 2/5 | -ve |
| 28 | 40 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 29 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 30 | 640 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 31 | 320 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 32 | 80 | 0 | 4/5 | 5/5 | 5/5 | +ve |
| 33 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 34 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 35 | 320 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 36 | 80 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 37 | 320 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 38 | 0 | 0 | 1/5 | 1/5 | 1/5 | -ve |
| 39 | 320 | 640 | 5/5 | 5/5 | 5/5 | +ve |
| 40 | 80 | 160 | 4/5 | 3/5 | 3/5 | +ve |
| 41 | 80 | 20 | 5/5 | 5/5 | 5/5 | +ve |
| 42 | 40 | 0 | 5/5 | 5/5 | 5/5 | +ve |
| 43 | 1280 | 320 | 4/5 | 1/5 | 1/5 | +ve |
| 44 | 0 | 0 | 1/5 | 2/5 | 2/5 | -ve |
| 45 | 0 |  | 4/5 | 5/5 | 4/5 | +ve |
| 46 | 0 | 160 | 5/5 | 5/5 | 4/5 | +ve |
| 47 | 40 | 80 | 3/5 | 2/5 | 2/5 | -ve |
| 48 | 0 | 160 | 4/5 | 5/5 | 4/5 | +ve |
| 49 | 320 | 160 | 5/5 | 5/5 | 5/5 | +ve |
| 50 | 160 | 160 | 5/5 | 5/5 | 5/5 | +ve |
| 51 | 5120 | 1280 | 5/5 | 4/5 | 4/5 | +ve |
| 52 | 640 | 160 | 5/5 | 5/5 | 5/5 | +ve |

Key:

1/5 = Negative (-ve)

2/5 = Inconclusive

3,4,5/5 = Positive (+ve)

“Leptorapide” Evaluation

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| "Leptorapide" Positive | 5 | 9 | 14 |
| "Leptorapide" Negative | 38 | 0 | 38 |
| Totals | 43 | 9 | 52 |

|  | Estimated Value | 95% Confidence Interval |  |
| --- | --- | --- | --- |
|  |  | Lower Limit | Upper Limit |
| Prevalence | 0.173077 | 0.0869 | 0.308194 |
| Sensitivity | 1 | 0.628811 | 1 |
| Specificity | 0.883721 | 0.741183 | 0.95639 |
| For any particular test result, the probability that it will be: |  |  |  |
| Positive | 0.269231 | 0.159978 | 0.412655 |
| Negative | 0.730769 | 0.587345 | 0.840022 |
| For any particular positive test result, the probability that it is: |  |  |  |
| True Positive | 0.642857 | 0.356284 | 0.860214 |
| False Positive | 0.357143 | 0.139786 | 0.643716 |
| For any particular negative test result, the probability that it is: |  |  |  |
| True Negative | 1 | 0.885666 | 1 |
| False Negative | 0 | 0 | 0.114334 |
| likelihood Ratios: |  |  |  |
| [C] = conventional |  |  |  |
| [W] = weighted by prevalence |  |  |  |
| Positive [C] | 8.6 | 3.772645 | 19.604284 |
| Negative [C] | 0 | 0 | NaN |
| Positive [W] | 1.8 | 0.805602 | 4.021835 |
| Negative [W] | 0 | 0 | NaN |

Appendix C – Combined Hereford UK and K.I.T Netherlands Calculations

“Leptorapide” Evaluation

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| “Leptorapide” Positive | 7 | 33 | 41 |
| “Leptorapide” Negative | 66 | 1 | 66 |
| Totals | 73 | 34 | 107 |

|  | Estimated Value |  | 95% Confidence Interval |  |
| --- | --- | --- | --- | --- |
|  |  |  | Lower Limit | Upper Limit |
| Prevalence | 0.317757 |  | 0.232986 | 0.415816 |
| Sensitivity | 0.970588 |  | 0.829462 | 0.998463 |
| Specificity | 0.90411 |  | 0.80667 | 0.957325 |
| For any particular test result, the probability that it will be: |  |  |  |  |
| Positive | 0.373832 |  | 0.283741 | 0.473106 |
| Negative | 0.626168 |  | 0.526894 | 0.716259 |
| For any particular positive test result, the probability that it is: |  |  |  |  |
| True Positive | 0.825 |  | 0.666382 | 0.921061 |
| False Positive | 0.175 |  | 0.078939 | 0.333618 |
| For any particular negative test result, the probability that it is: |  |  |  |  |
| True Negative | 0.985075 |  | 0.908613 | 0.99922 |
| False Negative | 0.014925 |  | 0.00078 | 0.091387 |
| likelihood Ratios: |  |  |  |  |
| [C] = conventional |  |  |  |  |
| [W] = weighted by prevalence |  |  |  |  |
| Positive [C] |  | 10.121849 | 4.992254 | 20.522159 |
| Negative [C] |  | 0.032531 | 0.004704 | 0.22497 |
| Positive [W] |  | 4.714286 | 2.369705 | 9.378589 |
| Negative [W] |  | 0.015152 | 0.002165 | 0.106041 |

Appendix D – Brazil Raw Data

| Sample Details |  | "Leptorapide" |  |  |  | Brazil |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Number | Spirochetes in Blood | Test 1 | Test 2 | Test 3 | Overall | IgM ELISA | Serovar Summary |
| 1888 | 1 | 5 | 5 | 5 | 2 | 2 | 400 |
| 282 | 2 | 4 | 4 | 5 | 2 | 2 | 100 |
| 123 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 161 | 1 | 5 | 5 | 5 | 2 | 2 | 400 |
| 223 | 2 | 5 | 5 | 5 | 2 | 2 | 100 |
| 265 | 1 | 4 | 4 | 5 | 2 | 2 | 100 |
| 282 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 297 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 319 | 2 | 5 | 5 | 5 | 2 | 2 | 100 |
| 370 | 1 | 5 | 5 | 5 | 2 | 2 | 400 |
| 375 | 2 | 5 | 5 | 5 | 2 | 2 | 100 |
| 401 | 2 | 2 | 5 | 5 | 2 | 2 | 100 |
| 415 | 2 | 5 | 5 | 5 | 2 | 2 | 100 |
| 451 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 471 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 483 | 2 | 3 | 5 | 5 | 2 | 2 | 200 |
| 589 | 2 | 3 | 4 | 5 | 2 | 2 | 200 |
| 630 | 1 | 5 | 5 | 5 | 2 | 2 | 3200 |
| 939 | 1 | 5 | 5 | 5 | 2 | 2 | 1600 |
| 1116 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 1150 | 2 | 4 | 5 | 5 | 2 | 2 | 100 |
| 1221 | 1 | 4 | 5 | 5 | 2 | 2 | 100 |
| 1421 | 1 | 5 | 5 | 5 | 2 | 2 | 200 |
| 2024 | 1 | 4 | 5 | 5 | 2 | 2 | 200 |
| 2082 | 2 | 5 | 5 | 5 | 2 | 2 | 100 |
| 2083 | 1 | 5 | 5 | 5 | 2 | 2 | 400 |
| 2102 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 2179 | 2 | 5 | 5 | 5 | 2 | 2 | 200 |
| 2187 | 2 | 5 | 5 | 5 | 2 | 2 | 400 |
| 2333 | 1 | 1 | 3 | 5 | 2 | 2 | 1600 |
| 2340 | 1 | 5 | 5 | 5 | 2 | 2 | 400 |
| 2371 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 2494 | 1 | 5 | 5 | 5 | 2 | 2 | 1600 |
| 2244 | 1 | 5 | 5 | 5 | 2 | 2 | 200 |
| 2316 | 1 | 5 | 5 | 5 | 2 | 2 | 1600 |
| 2321 | 1 | 3 | 5 | 5 | 2 | 2 | 200 |
| 2322 | 2 | 5 | 5 | 5 | 2 | 2 | 200 |
| 271 | 2 | 5 | 5 | 5 | 2 | 2 | 400 |
| 2151 | 1 | 4 | 4 | 5 | 2 | 2 | 200 |
| 2155 | 1 | 5 | 5 | 5 | 2 | 2 | 100 |
| 667 | 1 | 5 | 5 | 4 | 2 | 2 | 3200 |
| 676 | 1 | 1 | 2 | 4 | 1 | 2 | 200 |
| 710 | 1 | 4 | 3 | 4 | 2 | 2 | 100 |
| 2056 | 1 | 4 | 4 | 4 | 2 | 2 | 200 |
| 2183 | 1 | 4 | 5 | 4 | 2 | 2 | 400 |
| 2234 | 1 | 3 | 3 | 4 | 2 | 2 | 100 |
| 2182 | 1 | 1 | 3 | 4 | 2 | 2 | 100 |
| 1045 | 1 | 1 | 3 | 4 | 2 | 2 | 800 |
| 249 | 2 | 5 | 5 | 3 | 2 | 2 | 100 |
| 611 | 1 | 2 | 4 | 3 | 2 | 2 | 800 |
| 723 | 1 | 1 | 3 | 3 | 2 | 2 | 100 |
| 729 | 1 | 5 | 5 | 3 | 2 | 2 | 100 |
| 1351 | 1 | 1 | 3 | 3 | 2 | 2 | 100 |
| 2219 | 1 | 5 | 4 | 3 | 2 | 2 | 100 |
| 713 | 1 | 1 | 2 | 3 | 1 | 2 | 100 |
| 1342 | 1 | 1 | 2 | 3 | 1 | 2 | 100 |
| 2192 | 1 | 1 | 2 | 3 | 1 | 2 | 200 |
| 276 | 1 | 5 | 5 | 3 | 2 | 2 | 200 |
| 544 | 1 | 1 | 2 | 2 | 1 | 2 | 200 |
| 593 | 1 | 1 | 1 | 2 | 0 | 2 | 1600 |
| 490 | 2 | 1 | 2 | 2 | 1 | 2 | 100 |
| 498 | 2 | 1 | 1 | 2 | 0 | 2 | 200 |
| 2622 | 1 | 1 | 3 | 2 | 1 | 2 | 100 |
| 946 | 1 | 4 | 2 | 1 | 1 | 2 | 100 |
| 2370 | 1 | 3 | 4 | 1 | 2 | 2 | 100 |
| 413 | 2 | 1 | 2 | 1 | 0 | 2 | 200 |
| 139 | 1 | 5 | 5 | 5 | 2 | 1 | 200 |
| 149 | 2 | 5 | 5 | 5 | 2 | 1 | 200 |
| 367 | 1 | 5 | 5 | 5 | 2 | 1 | 400 |
| 385 | 1 | 5 | 5 | 5 | 2 | 1 | 400 |
| 424 | 1 | 5 | 5 | 5 | 2 | 1 | 100 |
| 622 | 2 | 5 | 5 | 5 | 2 | 1 | 100 |
| 742 | 1 | 5 | 5 | 5 | 2 | 1 | 400 |
| 1092 | 1 | 5 | 5 | 5 | 2 | 1 | 100 |
| 2102 | 1 | 1 | 2 | 5 | 1 | 1 | 100 |
| 2246 | 1 | 5 | 5 | 5 | 2 | 1 | 800 |
| 2262 | 1 | 5 | 5 | 5 | 2 | 1 | 400 |
| 832 | 1 | 5 | 5 | 4 | 2 | 1 | 200 |
| 869 | 1 | 1 | 3 | 4 | 2 | 1 | 200 |
| 509 | 2 | 1 | 1 | 3 | 0 | 1 | 100 |
| 2189 | 1 | 1 | 2 | 3 | 1 | 1 | 100 |
| 2241 | 1 | 2 | 2 | 3 | 1 | 1 | 100 |
| 2126 | 1 | 1 | 2 | 3 | 1 | 1 | 100 |
| 1053 | 1 | 2 | 3 | 2 | 1 | 1 | 400 |
| 2093 | 1 | 3 | 2 | 2 | 1 | 1 | 400 |
| 2116 | 1 | 1 | 1 | 2 | 0 | 1 | 800 |
| 2122 | 1 | 1 | 1 | 2 | 0 | 1 | 200 |
| 2176 | 1 | 1 | 2 | 2 | 1 | 1 | 200 |
| 2133 | 1 | 1 | 2 | 1 | 0 | 1 | 100 |
| 1033 | 1 | 1 | 2 | 1 | 0 | 1 | 200 |
| 423 | 1 | 2 | 3 | 5 | 2 | 0 | 200 |
| 467 | 1 | 5 | 5 | 5 | 2 | 0 | 400 |
| 2233 | 1 | 1 | 3 | 5 | 2 | 0 | 200 |
| 1051 | 1 | 2 | 3 | 5 | 2 | 0 | 400 |
| 1068 | 1 | 5 | 5 | 5 | 2 | 0 | 100 |
| 754 | 1 | 5 | 5 | 4 | 2 | 0 | 1600 |
| 2299 | 1 | 5 | 5 | 4 | 2 | 0 | 100 |
| 2105 | 1 | 1 | 3 | 4 | 2 | 0 | 400 |
| 2304 | 1 | 5 | 5 | 4 | 2 | 0 | 200 |
| 858 | 1 | 1 | 2 | 4 | 1 | 0 | 100 |
| 2084 | 1 | 4 | 2 | 4 | 2 | 0 | 100 |
| 2130 | 1 | 1 | 2 | 4 | 1 | 0 | 400 |
| 2156 | 1 | 1 | 2 | 4 | 1 | 0 | 100 |
| 2255 | 1 | 3 | 5 | 4 | 2 | 0 | 100 |
| 2264 | 1 | 1 | 3 | 4 | 2 | 0 | 100 |
| 681 | 1 | 1 | 2 | 3 | 1 | 0 | 200 |
| 711 | 1 | 1 | 2 | 3 | 1 | 0 | 100 |
| 922 | 1 | 5 | 5 | 3 | 2 | 0 | 100 |
| 2117 | 1 | 3 | 4 | 3 | 2 | 0 | 400 |
| 922 | 1 | 5 | 2 | 3 | 2 | 0 | 100 |
| 693 | 1 | 1 | 1 | 3 | 0 | 0 | 200 |
| 703 | 1 | 1 | 2 | 3 | 1 | 0 | 200 |
| 1072 | 1 | 3 | 1 | 3 | 2 | 0 | 100 |
| 2184 | 1 | 1 | 2 | 3 | 1 | 0 | 100 |
| 2274 | 1 | 1 | 1 | 3 | 0 | 0 | 100 |
| 2090 | 1 | 1 | 1 | 3 | 0 | 0 | 200 |
| 2119 | 1 | 3 | 4 | 3 | 2 | 0 | 200 |
| 2175 | 1 | 1 | 2 | 3 | 1 | 0 | 100 |
| 2202 | 2 | 1 | 1 | 3 | 0 | 0 | 100 |
| 2229 | 1 | 1 | 2 | 3 | 1 | 0 | 100 |
| 2257 | 1 | 2 | 2 | 3 | 1 | 0 | 100 |
| 366 | 1 | 1 | 2 | 2 | 1 | 0 | 100 |
| 2121 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 1073 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 716 | 1 | 1 | 2 | 2 | 1 | 0 | 100 |
| 920 | 1 | 2 | 2 | 2 | 1 | 0 | 100 |
| 2147 | 2 | 1 | 2 | 2 | 1 | 0 | 100 |
| 349 | 1 | 5 | 5 | 2 | 2 | 0 | 400 |
| 431 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 511 | 1 | 1 | 2 | 2 | 1 | 0 | 100 |
| 2092 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 2097 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 2098 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 2111 | 1 | 1 | 1 | 2 | 0 | 0 | 200 |
| 2118 | 1 | 1 | 2 | 2 | 1 | 0 | 100 |
| 2190 | 1 | 1 | 1 | 2 | 0 | 0 | 100 |
| 752 | 2 | 2 | 2 | 1 | 1 | 0 | 100 |
| 921 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 2415 | 1 | 1 | 1 | 1 | 0 | 0 | 200 |
| 2451 | 1 | 1 | 1 | 1 | 0 | 0 | 200 |
| 666 | 1 | 2 | 1 | 1 | 0 | 0 | 200 |
| 2137 | 1 | 2 | 2 | 1 | 1 | 0 | 100 |
| 2329 | 1 | 1 | 1 | 1 | 0 | 0 | 100 |
| 1008 | 1 | 1 | 3 | 1 | 0 | 0 | 100 |
| 1032 | 1 | 2 | 2 | 1 | 1 | 0 | 100 |
| 2534 | 1 | 1 | 1 | 1 | 0 | 0 | 100 |
| 670 | 1 | 1 | 1 | 1 | 0 | 0 | 100 |
| 704 | 2 | 1 | 1 | 1 | 0 | 0 | 100 |
| 2091 | 1 | 2 | 2 | 1 | 1 | 0 | 100 |
| 2157 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 2169 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 2171 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 1046 | 1 | 1 | 2 | 1 | 0 | 0 | 200 |
| 394 | 1 | 1 | 1 | 1 | 0 | 0 | 100 |
| 1419 | 1 | 2 | 2 | 1 | 1 | 0 | 100 |
| 2142 | 1 | 1 | 2 | 1 | 0 | 0 | 400 |
| 2143 | 1 | 1 | 2 | 1 | 0 | 0 | 200 |
| 2164 | 2 | 1 | 2 | 1 | 0 | 0 | 100 |
| 1009 | 1 | 2 | 1 | 1 | 0 | 0 | 200 |
| 1022 | 2 | 2 | 1 | 1 | 0 | 0 | 400 |
| 1030 | 1 | 1 | 2 | 1 | 0 | 0 | 800 |
| 1036 | 1 | 1 | 2 | 1 | 0 | 0 | 200 |
| 1037 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 1047 | 2 | 1 | 2 | 1 | 0 | 0 | 200 |
| 1050 | 1 | 1 | 2 | 1 | 0 | 0 | 100 |
| 1052 | 1 | 5 | 4 | 1 | 2 | 0 | 1600 |
| 1058 | 1 | 5 | 5 | 1 | 2 | 0 | 100 |
| 1059 | 1 | 1 | 1 | 1 | 0 | 0 | 100 |

Key: 					“Leptorapide” scoring;

0 = Negative				1/5 = Negative

1 = Inconclusive			2/5 = Inconclusive

2 = Positive			           3,4,5/5 = Positive

Appendix E – Poland Raw Data and Calculations

| Details |  | "Leptorapide" Results |  |  |
| --- | --- | --- | --- | --- |
| ID No | MAT | Test 1 | Test 2 | Overall |
| 1 | Neg | 1/5 | 1/5 | Neg |
| 2 | Neg | 2/5 | 3/5 | Positive |
| 3 | Neg | 1/5 | 1/5 | Neg |
| 4 | Neg | 1/5 | 2/5 | Neg |
| 5 | Neg | 1/5 | 1/5 | Neg |
| 6 | Neg | 1/5 | 1/5 | Neg |
| 7 | Neg | 1/5 | 1/5 | Neg |
| 8 | Neg | 1/5 | 1/5 | Neg |
| 9 | Neg | 1/5 | 1/5 | Neg |
| 10 | Neg | 1/5 | 1/5 | Neg |
| 11 | Neg | 1/5 | 1/5 | Neg |
| 12 | Neg | 1/5 | 1/5 | Neg |
| 13 | Neg | 1/5 | 1/5 | Neg |
| 14 | Neg | 1/5 | 1/5 | Neg |
| 15 | Neg | 1/5 | 1/5 | Neg |
| 16 | Neg | 1/5 | 1/5 | Neg |
| 17 | Neg | 1/5 | 1/5 | Neg |
| 18 | Neg | 1/5 | 1/5 | Neg |
| 19 | Neg | 1/5 | 1/5 | Neg |
| 20 | Neg | 3/5 | 4/5 | Positive |
| 21 | Neg | 1/5 | 1/5 | Neg |
| 22 | Neg | 1/5 | 1/5 | Neg |
| 23 | Neg | 2/5 | 3/5 | Positive |
| 24 | Neg | 1/5 | 2/5 | Neg |
| 25 | Neg | 1/5 | 1/5 | Neg |
| 26 | Neg | 1/5 | 1/5 | Neg |
| 27 | Neg | 5/5 | 5/5 | Positive |
| 28 | Neg | 2/5 | 2/5 | Neg |
| 29 | Neg | 1/5 | 1/5 | Neg |
| 30 | Neg | 1/5 | 1/5 | Neg |
| 31 | Neg | 1/5 | 1/5 | Neg |
| 32 | Neg | 1/5 | 1/5 | Neg |
| 33 | Neg | 1/5 | 1/5 | Neg |
| 34 | Neg | 1/5 | 1/5 | Neg |
| 35 | Neg | 1/5 | 1/5 | Neg |
| 36 | Neg | 1/5 | 1/5 | Neg |
| 37 | Neg | 1/5 | 1/5 | Neg |
| 38 | Neg | 1/5 | 1/5 | Neg |
| 39 | Neg | 1/5 | 1/5 | Neg |
| 40 | Neg | 1/5 | 1/5 | Neg |
| 41 | Neg | 1/5 | 1/5 | Neg |
| 42 | Neg | 1/5 | 1/5 | Neg |
| 43 | Neg | 1/5 | 1/5 | Neg |
| 44 | Neg | 1/5 | 1/5 | Neg |
| 45 | Neg | 1/5 | 1/5 | Neg |
| 46 | Neg | 1/5 | 1/5 | Neg |
| 47 | Neg | 1/5 | 1/5 | Neg |
| 48 | Neg | 1/5 | 1/5 | Neg |
| 49 | Neg | 1/5 | 1/5 | Neg |
| 50 | Neg | 1/5 | 1/5 | Neg |
| 51 | Neg | 1/5 | 1/5 | Neg |
| 52 | Neg | 2/5 | 2/5 | Neg |
| 53 | Neg | 1/5 | 1/5 | Neg |
| 54 | Neg | 1/5 | 1/5 | Neg |
| 55 | Neg | 2/5 | 2/5 | Neg |
| 56 | Neg | 1/5 | 1/5 | Neg |
| 57 | Neg | 1/5 | 1/5 | Neg |
| 58 | Neg | 1/5 | 2/5 | Neg |
| 59 | Neg | 1/5 | 1/5 | Neg |
| 60 | Neg | 1/5 | 1/5 | Neg |
| 61 | Neg | 1/5 | 1/5 | Neg |
| 62 | Neg | 1/5 | 1/5 | Neg |
| 63 | Neg | 1/5 | 1/5 | Neg |
| 64 | Neg | 1/5 | 1/5 | Neg |
| 65 | Neg | 1/5 | 1/5 | Neg |
| 66 | Neg | 1/5 | 1/5 | Neg |
| 67 | Neg | 1/5 | 1/5 | Neg |
| 68 | Neg | 1/5 | 1/5 | Neg |
| 69 | Neg | 1/5 | 1/5 | Neg |
| 70 | Neg | 1/5 | 1/5 | Neg |
| 71 | Neg | 1/5 | 1/5 | Neg |
| 72 | Neg | 1/5 | 1/5 | Neg |
| 73 | Neg | 1/5 | 1/5 | Neg |
| 74 | Neg | 1/5 | 1/5 | Neg |
| 75 | Neg | 1/5 | 1/5 | Neg |
| 76 | Neg | 1/5 | 1/5 | Neg |
| 77 | Neg | 1/5 | 1/5 | Neg |
| 78 | Neg | 1/5 | 1/5 | Neg |
| 79 | Neg | 1/5 | 1/5 | Neg |
| 80 | Neg | 1/5 | 1/5 | Neg |
| 81 | Neg | 1/5 | 1/5 | Neg |
| 82 | Neg | 1/5 | 1/5 | Neg |
| 83 | Neg | 1/5 | 1/5 | Neg |
| 84 | Neg | 1/5 | 1/5 | Neg |
| 85 | Neg | 1/5 | 1/5 | Neg |
| 86 | Neg | 1/5 | 1/5 | Neg |
| 87 | Neg | 1/5 | 1/5 | Neg |
| 88 | Neg | 1/5 | 1/5 | Neg |
| 89 | Neg | 1/5 | 1/5 | Neg |
| 90 | Neg | 1/5 | 1/5 | Neg |
| 91 | Neg | 1/5 | 1/5 | Neg |
| 92 | Neg | 1/5 | 1/5 | Neg |
| 93 | Neg | 1/5 | 1/5 | Neg |
| 94 | Neg | 1/5 | 1/5 | Neg |
| 95 | Neg | 1/5 | 1/5 | Neg |
| 96 | Neg | 1/5 | 1/5 | Neg |

“Leptorapide” scoring;

1/5 = Negative

2/5 = Inconclusive

3,4,5/5 = Positive

“Leptorapide” Evaluation

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| Test Positive | 4 | 0 | 4 |
| Test Negative | 92 | 0 | 92 |
| Totals | 96 | 0 | 96 |

|  | Estimated Value |  | 95% Confidence Interval |  |
| --- | --- | --- | --- | --- |
|  |  |  | Lower Limit | Upper Limit |
| Prevalence | 0 |  | 0 | 0.047947 |
| Sensitivity | NaN |  | NaN | NaN |
| Specificity | 0.958333 |  | 0.890714 | 0.986569 |
| For any particular test result, the probability that it will be: |  |  |  |  |
| Positive | 0.041667 |  | 0.013431 | 0.109286 |
| Negative | 0.958333 |  | 0.890714 | 0.986569 |
| For any particular positive test result, the probability that it is: |  |  |  |  |
| True Positive | 0 |  | 0 | 0.604226 |
| False Positive | 1 |  | 0.395774 | 1 |
| For any particular negative test result, the probability that it is: |  |  |  |  |
| True Negative |  | 1 | 0.950053 | 1 |
| False Negative |  | 0 | 0 | 0.049947 |
| likelihood Ratios: |  |  |  |  |
| [C] = conventional |  |  |  |  |
| [W] = weighted by prevalence |  |  |  |  |
| Positive [C] | NaN |  | NaN | NaN |
| Negative [C] | NaN |  | NaN | NaN |
| Positive [W] | 0 |  | 0 | NaN |
| Negative [W] | 0 |  | 0 | NaN |

Appendix F – Italy Raw Data and Calculations

| Details |  | Sample | "Leptorapide" Results |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Patient No. | MAT Titre | Week No. | Test 1 | Test 2 | Test 3 | Overall |
| 1 | 1:1000 | 1 | P3 | P3 | P5 | Pos |
| 2 | N | 1 | N | N | N | Neg |
| 3 | 1:200 | 1 | N | N | N | Neg |
| 4 | 1:1000 | 1 | P3 | P4 | P4 | Pos |
| 5 | N | 1 | N | N | N | Neg |
| 6 | N | 1 | N | N | N | Neg |
| 7 | N | 1 | N | N | N | Neg |
| 9 | N | 1 | N | N | N | Neg |
| 10 | N | 1 | N | N | N | Neg |
| 11 | N | 1 | N | N | N | Neg |
| 12 | 1:100 | 1 | N | N | N | Neg |
| 16 | N | 2 | N | N | N | Neg |
| 13/I | N | 1 | P4 | P4 | P4 | Pos |
| 13/II | 1:1600 | 2 | P3 |  |  | Pos |
| 13/III | 1:3200 | 3 | P4 | P3 | P3 | Pos |
| 14 | N | 1 | N | N | N | Neg |
| 15 | 1:1000 | 1 | P4 | P4 | P4 | Pos |
| 17 | 1:100 | 1 | N | N | N | Neg |
| 18 | N | 1 | N | N | N | Neg |
| 19 | N | 1 | N | N | N | Neg |
| 20 | N | 1 | N | N | N | Neg |
| 21 | N | 2 | N | N | N | Neg |
| 22 | N | 1 | N | N | N | Neg |
| 23 | N | 1 | N | N | N | Neg |
| 24 | N | 2 | N | N | N | Neg |
| 25 | 1:3200 | 1 | P4 | P4 | P4 | Pos |
| 27 | 1:3200 | 3 | P4 | P4 | P5 | Pos |
| 29 | N | 1 | N | N | N | Neg |
| 30 | N | 1 | N | N | N | Neg |
| 28 | N | 1 | N | N | N | Neg |
| 31 | N | 1 | N | N | N | Neg |
| 32 | N | 1 | N | N | N | Neg |
| 33 | N | 1 | N | N | N | Neg |
| 34 | N | 2 | N | N | N | Neg |
| 35 | N | 1 | N | N | N | Neg |
| 36 | 1:320 | 1 | N | N | N | Neg |
| 37 | N | 1 | N | N | N | Neg |
| 38 | N | 1 | N | N | N | Neg |
| 39 | N | 1 | N | N | N | Neg |
| 40 | N | 1 | N | N | N | Neg |
| 43 | N | 2 | P3 | P4 | D2/P3 | Pos |
| 41 | 1:3200 | 1 | P3 | P3 | P3 | Pos |
| 42 | N | 1 | N | N | N | Neg |
| 44 | N | 1 | N | N | N | Neg |
| 45 | N | 1 | P5 | P5 | P3 | Pos |
| 46 | N | 1 | N | N | N | Neg |
| 47 | 1:320 | 1 | P4 | P3/4 | P4 | Pos |
| 48 | N | 1 | N | N | N | Neg |
| 50 | 1:1600 | 1 | P3 | P3 | P5 | Pos |
| 51 | 1:100 | 1 | P3 | P3 | P3 | Pos |
| 52 | 1:100 | 1 | P4 | P4/5 | P3 | Pos |
| 54 | 1:3200 | 3 | P4 | P4/5 | P4 | Pos |
| 55 | N | 1 | P5 | P4/5 | P5 | Pos |
| 56 | 1:320 | 2 | P4 | P5 | P5 | Pos |
| 57 | 1:100 | 1 | N | N | N | Neg |
| 108 | 1:10000 | 1 | P3 | P4/5 | P5 | Pos |
| 109 | 1:10000 | 2 | P4 | P4 | P5 | Pos |
| 110 | 1:10000 | 1 | P3 | P4 | P4 | Pos |
| 111 | 1:10000 | 2 | P4 | P4/5 | P5 | Pos |
| 112 | N | 1 |  | N |  | Neg |
| 113 | 1:3200 | 1 | P4 | P5 | P5 | Pos |
| 114 | 1:100 | 1 | N | N | N | Neg |
| 115 | 1:1000 | 2 | N | N | N | Neg |
| 116 | N | 1 | N | N | N | Neg |
| 117 | N | 1 | P3 | P3 | P3 | Pos |
| 118 | N | 1 | N | N | N | Neg |
| 123 | N | 2 | N | N | N | Neg |
| 119 | N | 1 | N | N | N | Neg |
| 120 | N | 1 | N | N | N | Neg |
| 125 | N | 2 | N | N | N | Neg |
| 121 | N | 1 | N | N | N | Neg |
| 122 | N | 1 | P3 | P4/5 | P4/5 | Pos |
| 124 | N | 1 | N | N | N | Neg |
| 126 | N | 1 | N | N | N | Neg |
| 128 | N | 1 | N | N | N | Neg |
| 127 | N | 1 | N | N | N | Neg |
| 132 | ND | 1 | D2/P3 | P3 | P3 | Pos |
| 133 | 1:1000 | 1 | P3 | P4 | P3 | Pos |
| 134 | 1:1000 | 1 | P3 | P5 | P3 | Pos |
| 135 | 1:1000 | 2 | P4 | P3 | P4 | Pos |
| 137 | 1:3200 | 1 | P3 | P3 | P3 | Pos |
| 136 | N | 1 | N | N | N | Neg |
| 138 | N | 1 | N | N | N | Neg |
| 139 | N | 1 | N | N | N | Neg |
| 140 | N | 1 | N | N | N | Neg |
| 141 | N | 1 | N | N | N | Neg |
| 145 | N | 3 | N | N | N | Neg |
| 143 | N | 1 | N | N | N | Neg |
| 146 | N | 1 | N | N | N | Neg |
| 147 | 1:1000 | 1 | P3 | P3/4 | P4/5 | Pos |
| 148 | N | 1 | N | N | N | Neg |
| 149 | 1:320 | 1 | P4 | P4 | P4 | Pos |
| 150 | 1:320 | 2 | P3 | P4 | P3 | Pos |
| 151 | 1:1000 | 1 | P4 | P4 | P3 | Pos |
| 153 | 1:1000 | 2 | P4 | P4 | D2 | Pos |
| 152 | 1:10000 | 1 | N | N | N | Neg |
| 155 | 1:10000 | 2 | N | N | N | Neg |
| 154 | 1:100 | 1 | N | N | N | Neg |
| 165 | 1:100 | 2 | N | N | N | Neg |
| 185 | 1:100 WR | 3 | N | N | N | Neg |
| 156 | 1:100 | 1 | N | N | N | Neg |
| 157 | N | 1 | N | N | N | Neg |
| 158 | N | 1 | N | N | N | Neg |
| 159 | 1:100 | 1 | N | N | N | Neg |
| 160 | 1:3200 | 1 | P4 | P4 | P4 | Pos |
| 162 | 1:10000 | 2 | P3 | P3 | P3 | Pos |
| 163 | 1:10000 | 1 | P3 | P3 | P3 | Pos |
| 164 | 1:10000 | 2 | P3 | P3 | P3 | Pos |
| 166 | 1:100 | 1 | P3 | P3 | P3 | Pos |
| 167 | 1:100 | 2 | N | N | N | Neg |
| 175 | N | 1 | N | N | N | Neg |
| 176 | N | 2 | N | N | N | Neg |
| 177 | N | 1 | N | N | N | Neg |
| 178 | N | 2 | N | N | N | Neg |
| 179 | 1:200 | 1 | N | N | N | Neg |
| 180 | 1:100 | 2 | N | N | N | Neg |
| 186 | 1:400 | 1 | P3 | P3 | P3 | Pos |
| 187 | N | 1 | N | N | N | Neg |
| 188 | N | 1 | N | N | N | Neg |
| 189 | N | 1 | D2/P3 | P3 | P3 | Pos |
| 190 | 1:100 | 1 | P3/4 | P4 | P5 | Pos |
| 194 | N | 1 | N | N | N | Neg |
| 195 | N | 2 | P3 | P3 | P3 | Pos |
| 196 | N | 1 | N/D2 | N/D2 | N/D2 | Neg |
| 197 | N | 1 | N | N | N | Neg |
| 198 | N | 1 | N | N | N | Neg |
| 199 | 1:3200 | 1 | P3/4 | P5 | P3 | Pos |
| 200 | 1:1000 | 2 | P3/4 | P3/4 | P4 | Pos |
| 201 | 1:320 | 1 | N | N | N | Neg |
| 221 | N | 2 | N | N | N | Neg |
| 222 | N | 3 | N | N | N | Neg |
| 223 | 1:3200 | 1 | P4 | P4 | P4 | Pos |
| 224 | 1:3200 | 2 | P4 | P4 | P4 | Pos |
| 228 | 1:1600 | 1 | P4 | P4/5 | P3/4 | Pos |
| 229 | 1:1600 | 2 | P4 | P4 | P4 | Pos |
| 231 | N | 1 | N | N | N | Neg |
| 232 | 1:400 | 1 | P4 | P4 | P4 | Pos |
| 233 | 1:800 | 2 | P3 | P4/5 | D2 | Pos |
| 238 | N | 1 | P4 | P5 | P4 | Pos |
| 240 | 1:100 | 1 | P4 | P5 | P4 | Pos |
| 241 | N | 1 | N | N | N | Neg |
| 245 | N | 2 | N | N | N | Neg |
| 242 | 1:1600 | 1 | P3 | P3 | P3 | Pos |
| 243 | 1:3200 | 1 | P3 | P3 | P3 | Pos |
| 248 | 1:3200 | 2 | P4 | P4 | P4 | Pos |
| 244 | 1:100 | 1 | N | N | N | Neg |
| 247 | 1:1600 | 2 | P4 | P5 | P5 | Pos |
| 249 | N | 1 | N | N | N | Neg |
| 250 | N | 1 | N | N | N | Neg |
| 251 | N | 2 | N | N | N | Neg |
| 252 | 1:3200 | 1 | P4 | P4 | P4 | Pos |
| 253 | 1:800 | 2 | P3 | P3 | P3 | Pos |
| 258 | 1:6400 | 1 | D2/P3 | P4 | P4 | Pos |
| 279 | N | 1 | P4 | P4 | P4 | Pos |
| 281 | 1:12800 | 2 | P4 | P4 | P4 | Pos |
| 314 | 1:12800 | 1 | D2/P3 | P4 | P4 | Pos |
| 322 | 1:400 | 1 | P3 | P4 | P4 | Pos |
| 326 | 1:12800 | 1 | P3 | P4 | P4 | Pos |
| 333 | 1:1600 | 1 | P4 | P4 | P4 | Pos |
| 334 | 1:6400 | 2 | P5 | P5 | P4 | Pos |
| 328 | N | 1 | N | N | N | Neg |
| 339 | 1:200 | 1 | N | N | N | Neg |
| 338 | 1:1600 | 2 | N | N | N | Neg |
| 329 | N | 1 | N | N | N | Neg |
| 331 | N | 2 | P3 | P3 | P3 | Pos |
| 340 | 1:200 | 1 | P5 | P5 | P5 | Pos |
| 345 | 1:3200 | 2 | P4 | P5 | P4 | Pos |
| 346 | N | 1 | N | N | N | Neg |
| 347 | N | 2 | N | N | N | Neg |
| 377 | 1:12800 | 1 | P4 | P4 | P4 | Pos |
| 396 | 1:12800 | 1 | P3 | P4 | P4 | Pos |
| 397 | 1:25600 | 2 | P3 | P5 | P5 | Pos |
| 398 | 1:6400 | 1 | P4 | P4 | P4 | Pos |
| 406 | 1:3200 | 1 | P4/5 | P4 | P5 | Pos |
| 407 | 1:6400 | 2 | P3 | P3 | P4 | Pos |
| 425 | 1:6400 | 1 | P3/4 | P5 | P5 | Pos |
| 429 | 1:6400 | 1 | P3 | P4 | P4 | Pos |
| 447 | 1:1600 | 1 | P5 | P5 | P4 | Pos |
| 448 | 1:1600 | 2 | P4 | P4 | P4 | Pos |
| 462 | 1:12800 | 2 | P5 | P4 | P5 | Pos |
| 463 | N | 1 | D2 | N | D2 | Neg |
| 477 | 1:1600 | 1 | P4 | P4 | P4 | Pos |
| 480 | 1:800 | 1 | P4 | P4 | P5 | Pos |
| 486 | 1:6400 | 2 | P5 | P4 | P4/5 | Pos |
| 495 | 1:3200 | 3 | P3 | P5 | P5 | Pos |
| 487 | N | 1 | N | N | N | Neg |
| 488 | N | 1 | N | N | N | Neg |
| 504 | N | 1 | N | N | N | Neg |
| 507 | N | 2 | N | N | N | Neg |
| 505 | N | 1 | N | N | N | Neg |
| 508 | N | 2 | N | N | N | Neg |
| 506 | N | 1 | N | N | N | Neg |
| 509 | N | 1 | N | N | N | Neg |
| 510 | N | 1 | N | N | N | Neg |
| 511 | N | 1 | N | N | N | Neg |
| 511 | 1:3200 | 1 | P4 | P4 | P5 | Pos |
| 512 | N | 1 | N | N | N | Neg |
| 513 | N | 1 | N | N | N | Neg |
| 514 | N | 1 | N | N | N | Neg |
| 515 | 1:800 | 1 | P4 | P3 | P4 | Pos |
| 520 | 1:1600 | 1 | P4 | P4 | P5 | Pos |
| 527 | N | 1 | N | N | N | Neg |
| 533 | N | 2 | N | N | N | Neg |
| 529 | N | 1 | N | N | N | Neg |
| 534 | N | 2 | N | N | N | Neg |
| 535 | N | 1 | N | N | N | Neg |
| 536 | N | 1 | N | N | N | Neg |
| 537 | N | 1 | N | N | N | Neg |
| 538 | N | 1 | N | N | N | Neg |
| 539 | N | 1 | N | N | N | Neg |
| 541 | N | 1 | N | N | N | Neg |
| 542 | N | 1 | N | N | N | Neg |
| 543 | N | 1 | N | N | N | Neg |
| 544 | N | 1 | N | N | N | Neg |
| 545 | N | 2 | N | N | N | Neg |
| 546 | N | 1 | N | N | N | Neg |
| 547 | N | 1 | N | N | N | Neg |
| 548 | N | 1 | N | N | N | Neg |
| 550 | 1:6400 | 2 | P4 | P5 | P5 | Pos |
| 549 | N | 1 | N | N | N | Neg |

Key: 					“Leptorapide” scoring;

WR = Weakly Reactive		N = Negative

ND = Not Determined		D2 = Inconclusive

P3/P4/P5 = Positive

“Leptorapide” Evaluation – Overall

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| Test Positive | 11 | 78 | 89 |
| Test Negative | 110 | 21 | 131 |
| Totals | 121 | 99 | 220 |
|  | Estimated Value | 95% Confidence Interval |  |
|  |  | Lower Limit | Upper Limit |
| Prevalence | 0.45 | 0.383476 | 0.518299 |
| Sensitivity | 0.787879 | 0.691878 | 0.861062 |
| Specificity | 0.909091 | 0.839566 | 0.9515 |
| For any particular test result, the probability that it will be: |  |  |  |
| Positive | 0.404545 | 0.33969 | 0.472791 |
| Negative | 0.595455 | 0.527209 | 0.66031 |
| For any particular positive test result, the probability that it is: |  |  |  |
| True Positive | 0.876404 | 0.78551 | 0.93371 |
| False Positive | 0.123596 | 0.06629 | 0.21449 |
| For any particular negative test result, the probability that it is: |  |  |  |
| True Negative | 0.839695 | 0.763065 | 0.895824 |
| False Negative | 0.160305 | 0.104176 | 0.236935 |
| likelihood Ratios: |  |  |  |
| [C] = conventional |  |  |  |
| [W] = weighted by prevalence |  |  |  |
| Positive [C] | 8.666667 | 4.888277 | 15.365558 |
| Negative [C] | 0.233333 | 0.159367 | 0.341628 |
| Positive [W] | 7.090909 | 4.055656 | 12.397744 |
| Negative [W] | 0.190909 | 0.12864 | 0.283319 |

“Leptorapide” Evaluation – Week 1 Samples

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| Test Positive | 7 | 49 | 56 |
| Test Negative | 93 | 12 | 105 |
| Totals | 100 | 61 | 161 |
|  | Estimated Value | 95% Confidence Interval |  |
|  |  | Lower Limit | Upper Limit |
| Prevalence | 0.378882 | 0.30468 | 0.45896 |
| Sensitivity | 0.803279 | 0.677811 | 0.889996 |
| Specificity | 0.93 | 0.856234 | 0.96898 |
| For any particular test result, the probability that it will be: |  |  |  |
| Positive | 0.347826 | 0.275695 | 0.427345 |
| Negative | 0.652174 | 0.572655 | 0.724305 |
| For any particular positive test result, the probability that it is: |  |  |  |
| True Positive | 0.875 | 0.75314 | 0.944098 |
| False Positive | 0.125 | 0.055902 | 0.24686 |
| For any particular negative test result, the probability that it is: |  |  |  |
| True Negative | 0.885714 | 0.805199 | 0.936962 |
| False Negative | 0.114286 | 0.063038 | 0.194801 |
| likelihood Ratios: |  |  |  |
| [C] = conventional |  |  |  |
| [W] = weighted by prevalence |  |  |  |
| Positive [C] | 11.47541 | 5.557208 | 23.696258 |
| Negative [C] | 0.211528 | 0.127203 | 0.351753 |
| Positive [W] | 7 | 3.476147 | 14.096068 |
| Negative [W] | 0.129032 | 0.075579 | 0.220291 |

“Leptorapide” Evaluation – Week 2 and 3 Samples

|  | Leptospirosis |  |  |
| --- | --- | --- | --- |
|  | Absent | Present | Totals |
| Test Positive | 3 | 29 | 32 |
| Test Negative | 19 | 8 | 27 |
| Totals | 22 | 37 | 59 |
|  | Estimated Value | 95% Confidence Interval |  |
|  |  | Lower Limit | Upper Limit |
| Prevalence | 0.627119 | 0.491082 | 0.746582 |
| Sensitivity | 0.783784 | 0.613358 | 0.895758 |
| Specificity | 0.863636 | 0.64038 | 0.964102 |
| For any particular test result, the probability that it will be: |  |  |  |
| Positive | 0.542373 | 0.408472 | 0.670759 |
| Negative | 0.457627 | 0.329241 | 0.591528 |
| For any particular positive test result, the probability that it is: |  |  |  |
| True Positive | 0.90625 | 0.738335 | 0.97547 |
| False Positive | 0.09375 | 0.02453 | 0.261665 |
| For any particular negative test result, the probability that it is: |  |  |  |
| True Negative | 0.703704 | 0.496592 | 0.855024 |
| False Negative | 0.296296 | 0.144976 | 0.503408 |
| likelihood Ratios: |  |  |  |
| [C] = conventional |  |  |  |
| [W] = weighted by prevalence |  |  |  |
| Positive [C] | 5.747748 | 1.981143 | 16.675528 |
| Negative [C] | 0.250356 | 0.133665 | 0.468919 |
| Positive [W] | 9.666667 | 3.272975 | 28.550304 |
| Negative [W] | 0.421053 | 0.228354 | 0.776361 |