Supplementary Table S1. Details of all antigens and reference strains along with cut off values for each sample within the evaluation of Leptorapide

(a) England: A battery of antigens were used with reference strains taken from the following groups; Leptospira interrogans (serovars Australis, Autumnalis, Bataviae, Canicola, Hardjo, Hebdomadis, Icterohaemorrhagiae, Pomona, Pyrogenes, Saxkoebing), Leptospira borgpetersenii (serovars Ballum, Javanica, Mini, Sejroe, Tarassovi), Leptospira kirshneri (serovars Cypnoteri, Grippotyphosa), Leptospira weilii (serovar Celledoni) and Leptospira biflexa (serovar Patoc). Titres of 1:40 or above were considered positive.

(b) K.I.T Netherlands: The Leptospirosis Reference Centre, Netherlands used a serum starting dilution of 1:20, and two-fold serial dilutions with a panel of 16 serovars. The following antigens were used ; Leptospira interrogans (serovars Australis, Autumnalis, Bataviae, Canicola, Hardjo, Hebdomadis, Icterohaemorrhagiae, Pomona, Pyrogenes, Saxkoebing), Leptospira borgpetersenii (serovars Ballum, Javanica, Mini, Sejroe, Tarassovi), Leptospira kirshneri (serovars Cypnoteri, Grippotyphosa), Leptospira weilii (serovar Celledoni) and Leptospira biflexa (serovar Patoc). A serum titre of >1:160 was considered positive diagnostic criteria for the disease.

(c) Brazil: The Centre for Zoonoses used the following strains, with each antigen tested separately : Leptospira interrogans (serovars Australis,  Autumnalis, Bataviae, Bratislava, Canicola, Djasiman, Hardjo, Hebdomadis, Icterohaemorraghiae, Lai, Pomona,  Pyrogenes, Rachmati, Saxkoebing, Sentot, Wolffi),  Leptospira borgpetersenii (serovars Ballum,  Castellonis, Copenhageni, Hardjo, Javanica,  Mini, Poi, Sejroe, Tarassovi, Whitcombi,)  Leptospira kirshneri (serovars Butembo, Cypnoteri, Grippotyphosa,), Leptospira noguchii (serovars Louisiana, Panama, Proechimys),  Leptospira weilii (serovar Celledoni),  Leptospira meyeri (serovar Semaranga), Leptospira biflexa ( serovars Andamana, Patoc), Leptospira santarosai (serovars Brasilienis, Shermani) and Leptospira fainei (serovar Hurstbridge). A serum titre of > 1:100 was considered positive for serological diagnosis of leptospirosis.

(d) India: The following strains were used: Leptospira interrogans (serovars Australis, Canicola, Hardjo, Hebdomadis, Icterohaemorraghiae, Lai, Pomona, Pyrogenes, Rachmati ),  Leptospira borgpetersenii (serovars Ballum,  Poi, Tarassovi), Leptospira kirshneri (serovar Grippotyphosa ), Leptospira biflexa ( serovar Andamana). Criteria for serological diagnosis of leptospirosis using MAT testing was; a) sero-conversion from a titre of < 1 in 20 in the first sample and ≥ 1 in 80 in the second sample. b) four-fold rise in titre, a titre of 1 in 80 in the first sample, 1 in 320 or above in the second sample. c) A titre of 1 in 80 or 1 in 160 was considered as a significant titre in the case where second sample is not available. The optimal cut-off titre for single MAT was assessed based on plotting the ROC curve for different MAT titre obtained on confirmed cases and controls.

(e) Poland: The following strains were used: Leptospira interrogans (serovars Australis, Autumnalis, Bataviae, Canicola, Hebdomadis, Icterohaemorraghiae, Pomona, Pyrogenes), Leptospira borgpetersenii (serovars Ballum, Javanica, Mini, Sejroe, Tarassovi), Leptospira kirshneri (serovars Cypnoteri, Grippotyphosa), Leptospira weilii (serovar Celledoni).

(f) Italy: A panel of 20 antigens, representative of 14 serogroups and 19 serovars were used: Leptospira interrogans (serovars Australis, Autumnalis, Bratislava, Bataviae, Canicola, Hardjo, Hardjo Prajitno, Hebdomadis, Icterohaemorrhagiae, Pomona, Pyrogenes, Saxkoebing), Leptospira borgpetersenii (serovars Ballum, Javanica, Mini, Sejroe, Tarassovi, Copenhageni,), Leptospira kirshneri (serovars Grippotyphosa). A serum titre of > 1:100 was considered positive for serological diagnosis of leptospirosis.

(g) Atlanta: The following strains were used: Leptospira interrogans (serovars Australis, Autumnalis, Bataviae, Bratislava, Canicola, Djasiman, Icterohaemorraghiae, Pomona, Pyrogenes, Mankarso, Wolffi), Leptospira borgpetersenii (serovars Ballum, Copenhageni, Javanica, Tarassovi), Leptospira kirshneri (serovars Cypnoteri, Grippotyphosa), Leptospira weilii (serovar Celledoni), Leptospira santarosai (serovars Borincana, Georgia). A serum titre of > 1:100 was considered positive for serological diagnosis of leptospirosis.

Supplementary Table S2.  Sensitivity and Specificity values for “Leptorapide” in relation to the MAT, from Combined Hereford UK and K.I.T Netherlands Sera Samples

Serum samples collected from K.I.T Netherlands and Hereford UK were tested with “Leptorapide”. Of the 107 serum samples, 34 were MAT positive and 73 MAT negative. The extent of agglutination was scored as described in Figure 1. Sensitivity and specificity of “Leptorapide” was calculated in relation to the MAT. “Leptorapide” compared well to the MAT, with sensitivity and specificity values of 97.1% and 94.0% respectively. The accuracy was calculated at 92.5% and Kappa Index 98%, showing “Leptorapide” is in agreement with the MAT in this study.

Supplementary Table S3. Comparison of IgM ELISA with “Leptorapide” on 168 MAT positive (>1:100) sera samples collected from Brazil

168 samples with a MAT titre >1:100 were obtained from patients with suspected leptospirosis in Brazil. The commercially available PanBio IgM ELISA was used to differentiate acute, on-going infection from previous convalescent infections. The ELISA deemed 66/168 of the samples to contain significant levels of specific IgM. “Leptorapide” detected 55 of these samples; a further 3 deemed “Leptorapide” negative and 8 samples were inconclusive, requiring a repeat test. 37 of the MAT positive sera were deemed negative by “Leptorapide” and IgM ELISA. A further 20 of the MAT positive sera were positive by “Leptorapide” but negative by IgM ELISA. 3 of the sera samples were positive for the MAT and IgM ELISA, but negative for “Leptorapide”.

Supplementary Table S4.  Sensitivity and Specificity of 4 Rapid Tests from Samples Taken from South Andaman ≤ 1 Week from Onset of Illness

Samples were collected at time of reporting illness from a Primary Health Centre (PHC) from the endemic region of South Andaman, India. Rapid tests were performed as per manufacturer’s protocol. MAT was performed using a panel of 15 serovars. Each patient sample was evaluated and indices calculated against standard diagnostic criteria. “Leptorapide” was the only test evaluated that had a sensitivity above 50%, with Leptocheck being the next most accurate test with 50% sensitivity. Each test had respectable specificity values, with “Leptorapide” being 87.2% specific.

Supplementary Table S5.  Sensitivity and Specificity of 4 Rapid Tests from Samples Taken from South Andaman ≥ 1 Week from Onset of Illness

Samples were collected 7 – 10 days after reporting illness to a Primary Health Centre (PHC) from the endemic region of South Andaman, India. Rapid tests were performed as per manufacturer’s protocol. MAT was performed using a panel of 15 serovars. Each patient sample was evaluated and indices calculated against standard diagnostic criteria. “Leptorapide” was again the most sensitive test evaluated, with sensitivity and specificity values of 90.0% and 87.3% respectively. The LAT was the second most accurate test, with sensitivity of 73.8% and specificity of 88.4%.

Supplementary Table S6. Sensitivity and Specificity of Leptorapide with Samples from a Polish Leptospirosis Reference Centre.

96 samples collected form the reference centre were confirmed negative by MAT.  The samples were tested with Leptorapide, of the 96 samples “Leptorapide” confirmed 92 as negative and 4 as positive.  Results showing that “Leptorapide” compares well to MAT with a specificity of 95.8%.